CASE NARRATIVE

### Matrix New World Engineering

### Project Name: Wonder Group Inc. Emergency IH Assess - 25-0580

### Project # N/A

### Order ID # Q2750

Test Name: TO-15

A. Number of Samples and Date of Receipt:

10 Air sample was received on 08/01/2025

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

TO-15. This data package contains results for TO-15.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_L were done using GC column RTX-1, which is 60 meters, 0.32 mm id, 1.0 um df, Restek Cat. #10157. The Trap was supplied by Entech, glass bead and Tenax , Entech 7100A Preconcentrator.The analysis of TO-15 was based on method TO-15.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries were met for all analysis.

The Internal Standards Areas were met for all analysis.

The Retention Times were met for all analysis.

The RPD for {Q2750-01DUP} with File ID: VL042800.D met criteria except for Tetrachloroethene[28.6%] due to difference in results of original and DUP.

The RPD for {Q2750-04DUP} with File ID: VL042814.D met criteria except for 1,2,4-Trimethylbenzene[32.6%], Carbon disulfide[20.7%], Hexane[23.7%], Methyl methacrylate[200%], Styrene[200%] and Toluene[27.8%] due to difference in results of original and DUP.

The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.

The %RSD is greater than 30% in the Initial Calibration method (VL072325AIR.M) for Naphthalene and 1,2,4-Trichlorobenzene these compounds are passing on Linear Regression.

The Continuous Calibration met requirements.

The Tuning criteria met requirements.

Samples WONDER IA-2, WONDER IA-3, WONDER IA-4, WONDER IA-5, WONDER IA-6, WONDER IA-7, WONDER IA-8, WONDER IA-9 and WONDER IA-11 were diluted due to high concentrations.

E. Additional Comments:

All Samples have the concentration of some target compound below Method detection limits, therefore it is not reported as Hit in Form1.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

The Manual Integrations are performed for the followings.

| Manual Integration Report |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sequence | VL072325 | Instrument | MSVOA_l |  |  |  |  |

| Sample ID | File ID | Parameter | Review By | Review On | Supervised By | Supervised On | Reason |
| --- | --- | --- | --- | --- | --- | --- | --- |
| VSTDICCC010 | VL042747.D | m/p-Xylene | SAM | 7/24/2025 9:02:56 AM | MMDadoda | 7/24/2025 2:20:57 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | 1,1,2-Trichloroethane | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | 1,4-Dioxane | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | 4-Methyl-2-Pentanone | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | cis-1,3-Dichloropropene | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | Cyclohexane | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | Heptane | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | m/p-Xylene | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | t-1,3-Dichloropropene | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC002 | VL042748.D | trans-1,2-Dichloroethene | SAM | 7/24/2025 9:03:01 AM | MMDadoda | 7/24/2025 2:20:59 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | 1,4-Dioxane | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | 4-Methyl-2-Pentanone | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | Bromodichloromethane | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | cis-1,3-Dichloropropene | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | Ethanol | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | Heptane | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | m/p-Xylene | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | Methyl Methacrylate | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC001 | VL042749.D | t-1,3-Dichloropropene | SAM | 7/24/2025 9:03:21 AM | MMDadoda | 7/24/2025 2:21:02 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | 1,1,2-Trichloroethane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | 1,2-Dibromoethane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | 1,4-Dioxane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | 2-Hexanone | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | 4-Methyl-2-Pentanone | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Benzyl Chloride | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Bromodichloromethane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | cis-1,3-Dichloropropene | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Dibromochloromethane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Ethanol | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Heptane | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | m/p-Xylene | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Methyl Methacrylate | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | t-1,3-Dichloropropene | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Tetrahydrofuran | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.5 | VL042750.D | Trichloroethene | SAM | 7/24/2025 9:03:31 AM | MMDadoda | 7/24/2025 2:21:04 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042752.D | 1,1,1-Trichloroethane | SAM | 7/24/2025 9:03:05 AM | MMDadoda | 7/24/2025 2:21:09 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042752.D | 1,1,2,2-Tetrachloroethane | SAM | 7/24/2025 9:03:05 AM | MMDadoda | 7/24/2025 2:21:09 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042752.D | Carbon Tetrachloride | SAM | 7/24/2025 9:03:05 AM | MMDadoda | 7/24/2025 2:21:09 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042752.D | Tetrachloroethene | SAM | 7/24/2025 9:03:05 AM | MMDadoda | 7/24/2025 2:21:09 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.03 | VL042752.D | Trichloroethene | SAM | 7/24/2025 9:03:05 AM | MMDadoda | 7/24/2025 2:21:09 PM | Peak Integrated by Software incorrectly |
| VSTDICC015 | VL042753.D | m/p-Xylene | SAM | 7/24/2025 9:03:41 AM | MMDadoda | 7/24/2025 2:21:10 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042754.D | 1,1,2,2-Tetrachloroethane | SAM | 7/24/2025 9:03:10 AM | MMDadoda | 7/24/2025 2:21:12 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042754.D | 1,2-Dibromoethane | SAM | 7/24/2025 9:03:10 AM | MMDadoda | 7/24/2025 2:21:12 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042754.D | Carbon Tetrachloride | SAM | 7/24/2025 9:03:10 AM | MMDadoda | 7/24/2025 2:21:12 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042754.D | Tetrachloroethene | SAM | 7/24/2025 9:03:10 AM | MMDadoda | 7/24/2025 2:21:12 PM | Peak Integrated by Software incorrectly |
| VSTDICC0.1 | VL042754.D | Trichloroethene | SAM | 7/24/2025 9:03:10 AM | MMDadoda | 7/24/2025 2:21:12 PM | Peak Integrated by Software incorrectly |
| VSTDICV010 | VL042755.D | 1,1,2-Trichloroethane | SAM | 7/24/2025 9:03:15 AM | MMDadoda | 7/24/2025 2:21:15 PM | Peak Integrated by Software incorrectly |
| VSTDICV010 | VL042755.D | m/p-Xylene | SAM | 7/24/2025 9:03:15 AM | MMDadoda | 7/24/2025 2:21:15 PM | Peak Integrated by Software incorrectly |

| Sequence | VL080425 | Instrument | MSVOA_l |
| --- | --- | --- | --- |

| Sample ID | File ID | Parameter | Review By | Review On | Supervised By | Supervised On | Reason |
| --- | --- | --- | --- | --- | --- | --- | --- |
| VSTDCCC010 | VL042796.D | m/p-Xylene | SAM | 8/5/2025 7:22:52 AM | MMDadoda | 8/5/2025 12:51:57 PM | Peak Integrated by Software incorrectly |
| VL0804ABS01 | VL042798.D | 1,1,2-Trichloroethane | SAM | 8/5/2025 7:22:57 AM | MMDadoda | 8/5/2025 12:52:02 PM | Peak Integrated by Software incorrectly |
| VL0804ABS01 | VL042798.D | m/p-Xylene | SAM | 8/5/2025 7:22:57 AM | MMDadoda | 8/5/2025 12:52:02 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | 2,2,4-Trimethylpentane | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Carbon Tetrachloride | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Heptane | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Isopropyl Alcohol | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | m/p-Xylene | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Propene | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Styrene | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Tetrachloroethene | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01 | VL042799.D | Tetrahydrofuran | SAM | 8/5/2025 7:23:02 AM | MMDadoda | 8/5/2025 12:52:05 PM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | 1,2,4-Trichlorobenzene | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | 2,2,4-Trimethylpentane | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Ethyl Acetate | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Heptane | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Isopropyl Alcohol | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | m/p-Xylene | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Propene | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Tetrachloroethene | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-01DUP | VL042800.D | Tetrahydrofuran | sam | 8/6/2025 2:03:17 AM | MMdadoda | 8/6/2025 2:04:31 AM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | 2,2,4-Trimethylpentane | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | Carbon Tetrachloride | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | Dichlorodifluoromethane | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | m/p-Xylene | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | Methyl Methacrylate | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | Propene | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-02 | VL042801.D | Toluene | SAM | 8/5/2025 7:24:16 AM | MMDadoda | 8/5/2025 12:52:12 PM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | 2,2,4-Trimethylpentane | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Ethyl Acetate | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Isopropyl Alcohol | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | m/p-Xylene | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Methyl Methacrylate | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Propene | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Styrene | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-03 | VL042802.D | Tetrahydrofuran | sam | 8/6/2025 2:03:22 AM | MMdadoda | 8/6/2025 2:04:36 AM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Carbon Tetrachloride | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Dichlorodifluoromethane | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Isopropyl Alcohol | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | m/p-Xylene | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Methyl Methacrylate | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Propene | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-04 | VL042803.D | Styrene | SAM | 8/5/2025 7:25:37 AM | MMDadoda | 8/5/2025 12:52:19 PM | Peak Integrated by Software incorrectly |
| Q2750-05 | VL042804.D | Carbon Tetrachloride | SAM | 8/5/2025 7:23:07 AM | MMDadoda | 8/5/2025 12:52:22 PM | Peak Integrated by Software incorrectly |
| Q2750-05 | VL042804.D | m/p-Xylene | SAM | 8/5/2025 7:23:07 AM | MMDadoda | 8/5/2025 12:52:22 PM | Peak Integrated by Software incorrectly |
| Q2750-05 | VL042804.D | Methyl Methacrylate | SAM | 8/5/2025 7:23:07 AM | MMDadoda | 8/5/2025 12:52:22 PM | Peak Integrated by Software incorrectly |
| Q2750-05 | VL042804.D | Propene | SAM | 8/5/2025 7:23:07 AM | MMDadoda | 8/5/2025 12:52:22 PM | Peak Integrated by Software incorrectly |
| Q2750-05 | VL042804.D | Toluene | SAM | 8/5/2025 7:23:07 AM | MMDadoda | 8/5/2025 12:52:22 PM | Peak Integrated by Software incorrectly |
| Q2750-06 | VL042805.D | 2,2,4-Trimethylpentane | SAM | 8/5/2025 7:25:03 AM | MMDadoda | 8/5/2025 12:52:25 PM | Peak Integrated by Software incorrectly |
| Q2750-06 | VL042805.D | Carbon Tetrachloride | SAM | 8/5/2025 7:25:03 AM | MMDadoda | 8/5/2025 12:52:25 PM | Peak Integrated by Software incorrectly |
| Q2750-06 | VL042805.D | Isopropyl Alcohol | SAM | 8/5/2025 7:25:03 AM | MMDadoda | 8/5/2025 12:52:25 PM | Peak Integrated by Software incorrectly |
| Q2750-06 | VL042805.D | Propene | SAM | 8/5/2025 7:25:03 AM | MMDadoda | 8/5/2025 12:52:25 PM | Peak Integrated by Software incorrectly |
| Q2750-06 | VL042805.D | Trichlorofluoromethane | SAM | 8/5/2025 7:25:03 AM | MMDadoda | 8/5/2025 12:52:25 PM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | 2,2,4-Trimethylpentane | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | Ethyl Acetate | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | Heptane | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | Isopropyl Alcohol | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | m/p-Xylene | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-07 | VL042806.D | Propene | sam | 8/6/2025 2:03:27 AM | MMdadoda | 8/6/2025 2:04:40 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | 2,2,4-Trimethylpentane | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Dichlorodifluoromethane | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Ethyl Acetate | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Heptane | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | m/p-Xylene | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Propene | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-08 | VL042807.D | Tetrahydrofuran | sam | 8/6/2025 2:03:34 AM | MMdadoda | 8/6/2025 2:04:45 AM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | 2,2,4-Trimethylpentane | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | Carbon Tetrachloride | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | Heptane | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | m/p-Xylene | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | Propene | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-09 | VL042808.D | Tetrahydrofuran | SAM | 8/5/2025 7:24:21 AM | MMDadoda | 8/5/2025 12:52:34 PM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | m/p-Xylene | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | Propene | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | tert-Butyl alcohol | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | Tetrahydrofuran | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10 | VL042809.D | Toluene | sam | 8/6/2025 2:03:40 AM | MMdadoda | 8/6/2025 2:04:50 AM | Peak Integrated by Software incorrectly |
| Q2750-10DL | VL042810.D | Propene | SAM | 8/5/2025 7:23:58 AM | MMDadoda | 8/5/2025 12:52:39 PM | Peak Integrated by Software incorrectly |
| Q2750-03DL | VL042813.D | Propene | SAM | 8/5/2025 7:23:12 AM | MMDadoda | 8/5/2025 12:52:44 PM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | 1,2,4-Trimethylbenzene | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | Carbon Tetrachloride | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | m/p-Xylene | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | Methyl Methacrylate | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | Propene | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DUP | VL042814.D | Styrene | sam | 8/6/2025 2:03:45 AM | MMdadoda | 8/6/2025 2:04:55 AM | Peak Integrated by Software incorrectly |
| Q2750-04DL | VL042815.D | Propene | SAM | 8/5/2025 7:24:42 AM | MMDadoda | 8/5/2025 12:52:50 PM | Peak Integrated by Software incorrectly |
| Q2750-05DL | VL042816.D | Propene | SAM | 8/5/2025 7:25:17 AM | MMDadoda | 8/5/2025 12:52:52 PM | Peak Integrated by Software incorrectly |
| Q2750-06DL | VL042817.D | Propene | SAM | 8/5/2025 7:23:18 AM | MMDadoda | 8/5/2025 12:52:53 PM | Peak Integrated by Software incorrectly |
| Q2750-07DL | VL042818.D | Propene | SAM | 8/5/2025 7:24:46 AM | MMDadoda | 8/5/2025 12:52:55 PM | Peak Integrated by Software incorrectly |
| Q2750-08DL | VL042819.D | Propene | SAM | 8/5/2025 7:23:40 AM | MMDadoda | 8/5/2025 12:52:57 PM | Peak Integrated by Software incorrectly |
| Q2750-01DL | VL042821.D | Propene | SAM | 8/5/2025 7:23:28 AM | MMDadoda | 8/5/2025 12:53:04 PM | Peak Integrated by Software incorrectly |
| Q2750-02DL | VL042822.D | Propene | SAM | 8/5/2025 7:23:35 AM | MMDadoda | 8/5/2025 12:53:06 PM | Peak Integrated by Software incorrectly |

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________