CASE NARRATIVE

### Sciacca General Contractors, LLC

### Project Name: 92 Clifford Street, Newark

### Project # N/A

### Order ID # Q2758

Test Name: SVOC-PAH

A.Number of Samples and Date of Receipt:

1 Solid sample was received on 08/01/2025.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

SVOC-PAH. This data package contains results for SVOC-PAH.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The analysis of SVOC-PAH was based on method 8270E and extraction was done based on method 3541.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries were met for all analysis.

The Internal Standards Areas were met for all analysis.

The Retention Times were met for all analysis.

The MS {Q2763-01MS} with File ID: BF143327.D recoveries met the requirements for all compounds except for Fluoranthene[36%], Recovery failed due to matrix interference, Therefor no further corrective action was taken.

The MSD recoveries met the requirements for all compounds.

The RPD for {Q2763-01MSD} with File ID: BF143328.D met criteria except for Benzo(a)anthracene[26%], Benzo(b)fluoranthene[28%], Benzo(g,h,i)perylene[36%], Fluoranthene[52%] and Phenanthrene[24%], RPD failed due to result difference between MS and MSD. Therefor no further corrective action was taken.

The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.

The Tuning criteria met requirements.

E. Additional Comments:

The soil samples results are based on a dry weight basis.

The Date and time of sampling was not listed in the COC.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________