| ORDER ID: Q2814 | MATRIX: Water |
| --- | --- |
| METHOD: 6010D,7470A |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution met criteria for all compounds. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (TW-WTS-13MS) analysis met criteria for all compounds except for Barium, Beryllium and Zinc due to Chemical interference during Digestion Process. The Matrix Spike Duplicate (TW-WTS-13MSD) analysis met criteria for all compounds except for Antimony, Beryllium and Zinc due to Chemical interference during Digestion Process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The Post Digest Spike (TW-WTS-13A) analysis met criteria for all compounds except for Barium, Beryllium and Zinc due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

Sample Q2814-17 had oily matrix, so to avoid damage to instrument sample is run with straight 5X dilution.

Sample Q2814-01, Q2814-03, Q2814-16, Q2814-17 analyzed as Total Metal and Sample Q2814-22, Q2814-23 analyzed as Dissolved Metal.

_______________________________________				__________________________

QA REVIEW								Date