[image: image1.png]AII|anoe

ECHNICA





284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
CASE NARRATIVE
Homeowner
Project Name: 178 Main St Franklin NJ - Lead Paint
Project # N/A
Order ID # Q2817
Test Name: Metals Group3
A. Number of Samples and Date of Receipt:

1 Solid samples were received on 08/11/2025.

B. Parameters:


According to the Chain of Custody document, the following analyses were requested:

Metals Group3. This data package contains results for Metals Group3.

C. Analytical Techniques:

The analysis of Metals Group3 was based on method 6010D and digestion based on method 3050 (soils).
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Blank Spike met requirements for all compounds.
The Duplicate analysis met criteria for all compounds.
The Matrix Spike analysis met criteria for all compounds.
The Matrix Spike Duplicate analysis met criteria for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Calibration met the requirements.      

The Serial Dilution met the acceptable requirements.
E. Additional Comments:
The Date and time of sampling was not listed in the COC.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


