| ORDER ID: Q2818 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% for certain compounds in the Initial Calibration (Method 8270-BF081225.M) for 2-Nitrophenol, Hexachlorocyclopentadiene, 2-Nitroaniline, 2,6-Dinitrotoluene, 4-Nitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylephenol, Pentachlorophenol, Butylbenzylphthalate, Bis(2-ethylhexylephthalate), Di-n-octyl phthalate. These Compounds is passing on Linear Regression . 2,4-Dinitrophenol are passing on Quadratic Regression. The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP0805.M) for 2-Nitrophenol , 2-Nitroaniline, 2,4-Dinitrotoluene, these Compounds are passing on Linear regression, and 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol these compounds are passing on Quadratic regression. The Continuous Calibration File ID BF143365.D met the requirements except for Bis(2-ethylhexyl)phthalate,Butylbenzylphthalate,Di-n-octyl phthalate and Pentachlorophenol but The associated samples have no positive hit of these compounds , therefore no corrective action was taken . The Continuous Calibration File ID BP025360.D met the requirements except for 2,4-Dinitrophenol,2-Nitrophenol,4,6-Dinitro-2-methylphenol and 2,4,6-Tribromophenol, The associated samples have no positive hit of these compounds , therefore no corrective action was taken . The Continuous Calibration File ID BP025377.D met the requirements except for 2,2-oxybis(1-Chloropropane),2,4-Dinitrophenol,2,6-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Pentachlorophenol and 2,4,6-Tribromophenol, The associated samples have no positive hit of these compounds , therefore no corrective action was taken . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q2830-05MS} with File ID: BF143367.D recoveries met the requirements for all compounds except for Butylbenzylphthalate[150%], Di-n-butylphthalate[133%], due to matrixinterference, therefore no corrective action is required due to matrix interference. The MSD {Q2830-05MSD} with File ID: BF143368.D recoveries met the acceptable requirements except for Butylbenzylphthalate[192%], Di-n-butylphthalate[142%], due to matrix interference, therefore no corrective action is required due to matrix interference. The Blank Spike for {PB169210BS} with File ID: BP025402.D met requirements for all samples except for 2,4,5-Trichlorophenol[100%], The associated samples have no positive hit of these compounds , therefore no corrective action was taken . . The RPD for {Q2830-05MSD} with File ID: BF143368.D met criteria except for 4-Nitroaniline[22%], 4-Nitrophenol[23%], Butylbenzylphthalate[25%], Caprolactam[22%] and Pyrene[31%], due to difference in results of MS and MSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date