| ORDER ID: Q2819 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration (Method 8270-BF082025.M) for Hexachlorocyclopentadiene, 2,4-Dinitrophenol these Compounds are passing on Linear regression. The Continuous Calibration File ID BF143539.D met the requirements except for Pentachlorophenol but no positive hit in associated sample therefore no corrective action taken. The Continuous Calibration File ID BP025487.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol but no positive hit in associated sample therefore no corrective action taken. The Continuous Calibration File ID BP025528.D met the requirements except for Benzaldehyde but no positive hit in associated samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries were met for all analysis except for 22BP-E [Terphenyl-d14 - 52%], 11M-S [Terphenyl-d14 - 48%], 11M-N [Terphenyl-d14 - 50%], 84SB-E [Terphenyl-d14 - 51%], 84SB-W [Terphenyl-d14 - 37%], 17M-S [Terphenyl-d14 - 52%], 17M-E [Terphenyl-d14 - 40%], 17M-W [Terphenyl-d14 - 51%] and 17M-N [Terphenyl-d14 - 47%] as per method one surrogate is allowed to failed, therefore no corrective action was taken also 22BP-S [2-Fluorobiphenyl - 37%, Terphenyl-d14 - 35%], 22BP-SRE [2,4,6-Tribromophenol - 38%, Terphenyl-d14 - 32%],84SB-S [2,4,6-Tribromophenol - 34%, 2-Fluorobiphenyl - 34%, Nitrobenzene-d5 - 36%, Terphenyl-d14 - 31%], 84SB-SRE [2,4,6-Tribromophenol - 32%, 2-Fluorobiphenyl - 43%, Terphenyl-d14 - 33%] All the failure samples in surrogates were reanalyzed to confirm the results as per method and reported in the data. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB169242BS} with File ID: BP025489.D met requirements for all compounds except for 1,4-Dioxane[65%] is failing marginally low therefore no corrective action taken. The RPD for {Q2819-17MSD} with File ID: BP025521.D met criteria except for 4-Chloroaniline[21%] due to difference in results of MS and MSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples 38M-S was diluted due to dirty matrix.

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

The not QT review data is reported in the Miscellaneous.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date