| ORDER ID: Q2832 | MATRIX: Solid |
| --- | --- |
| METHOD: 6010D,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution (0804-SOILL) met criteria for all samples except for Mercury due to sample matrix interference. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (0804-SOILMS) analysis met criteria for all samples except for Mercury due to sample Matrix interference. The Matrix Spike (TG-S05MS) analysis met criteria for all samples except for Chromium and Vanadium due to Chemical Interference during Digestion process. The Matrix Spike Duplicate (TG-S05MSD) analysis met criteria for all compounds except for Chromium and Sodium due to Chemical Interference during Digestion process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (0804-SOILDUP) analysis met criteria for all compounds except for Mercury due to sample matrix interference. The Duplicate (0804-SOILMSD) analysis met criteria for all compounds except for Mercury due to sample matrix interference. The Duplicate (TG-S05MSD) analysis met criteria for all compounds except for Manganese due to Chemical Interference during Digestion process. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS: The Post Digest Spike (TG-S05A) analysis met criteria for all compounds except for Chromium, Sodium and Vanadium due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

_______________________________________				__________________________

QA REVIEW								Date