[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

ENTACT
Project Name: 540 Degraw St, Brooklyn, NY - E9309
Project # N/A
Order ID # Q2837
Test Name: SVOCMS Group4
A. Number of Samples and Date of Receipt:

1 Water sample was received on 08/12/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOCMS Group4. This data package contains results for SVOCMS Group4.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. BNA_P using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID,0.5 um df, Catalog # 7HG-G027-17-GGA.The analysis of SVOCMS Group4 was based on method 8270E and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.

The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the Requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


