ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2842

	MATRIX: Water

	METHOD: 6020B,7470A

	


	
	
	NA
	NO
	YES

	   1.
	Calibration Summary met criteria.


	
	
	(

	   2.
	ICP Interference Check Sample Results Summary Submitted.


	
	
	(

	   3.
	Serial Dilution Summary (if applicable) Submitted.


	
	
	(

	   4.
	Laboratory Control Sample Summary (if applicable) Submitted.


	
	
	(

	   5.
	Blank Contamination - If yes, list compounds and concentrations in each blank:



	
	(
	

	   6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Matrix Spike (MW-17B-55.5-08122025MS) analysis met criteria for all compounds except for Silver due to Chemical Interference during Digestion Process.  The Matrix Spike Duplicate (MW-17B-55.5-08122025MSD) analysis met criteria for all compounds except for Silver due to Chemical Interference during Digestion Process.

	
	(
	

	   7.
	Sample Duplicate Analysis Met QC Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	   8.
	Digestion Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	   9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS:
             This Data Package has been revised due to Data Package type Change as per Client Request.
 The Post Digest Spike (MW-17B-55.5-08122025A) analysis met criteria for all compounds except for Silver due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Tl, Pb, U are being analyzed with Non-Collision Cell. Helium gas is used for the Collision Cell analysis.
Sample Q2842-03 analyzed as Total Metal and Sample Q2842-09 analyzed as Dissolved Metal.
_______________________________________



__________________________

QA REVIEW







Date


