284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922

CASE NARRATIVE

### Tully Environmental, Inc

### Project Name: Transfer Station-SPDES

### Project # N/A

### Order ID # Q2845

Test Name: Metals Group 10

A. Number of Samples and Date of Receipt:

2 Water samples were received on 08/13/2025.

B. Parameters:

According to the Chain of Custody document, the following analyses were requested:

Metals Group 10, Oil and Grease, TSS and VOC-BTEX. This data package contains results for Metals Group 10.

C. Analytical Techniques:

The analysis and digestion of Metals Group 10 was based on method 200.7.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Blank Spike met requirements for all compounds.

The Duplicate (002 35th Ave (july)MSD) analysis met criteria for all compounds except for Iron due to Chemical Interference during digestion Process.

The Matrix Spike (002 35th Ave (july)MS) analysis met criteria for all compounds except for Iron due to Chemical Interference during digestion Process.

The Matrix Spike Duplicate (002 35th Ave (july)MSD) analysis met criteria for all compounds except for Iron due to Chemical Interference during digestion Process.

The Blank analysis did not indicate the presence of lab contamination.

The Calibration met the requirements.

The Serial Dilution met the acceptable requirements.

E. Additional Comments:

The Post Digest Spike (002 35th Ave (july)A) analysis met criteria for all compounds except for  Iron due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.

The temperature of the samples at the time of receipt was 16.7°C.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________