[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Nobis Group
Project Name: Raymark Superfund Site
Project # N/A
Order ID # Q2879
Test Name: VOCMS Group3
A. Number of Samples and Date of Receipt:

9 Solid samples were received on 08/15/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group3. This data package contains results for VOCMS Group3.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group3 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for OU4-TS-GRILLO-TSCP11-081425 [4-Bromofluorobenzene - 76%], OU4-TS-GRILLO-TSCP11-081425RE [4-Bromofluorobenzene - 76%] sample was reanalyzed to confirm the failure and reported.
The Internal Standards Areas were met for all analysis except for OU4-TS-GRILLO-TSCP11-081425RE, sample was reanalyzed to confirm the failure and reported.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Initial Calibration method (82Y081225S.M) for Methylene chloride passing on Linear regression.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
Alliance has analyzed samples for VOCMS Group3 by Method 8260D for Project

“Raymark Superfund Site”. Alliance is not certified for trans-1,4-dichloro-2-butene and

Tetrahydrofuran compounds with NJDEP for 8260D method. for VOCMS Group3 at the

time when samples for Project “Raymark Superfund Site “were analyzed.

As per special requirement for this project form-1 are reported in mg/kg.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The not QT review data is reported in the Miscellaneous.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


