ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2885

	MATRIX: Water

	METHOD: 8270-Modified/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The Initial Calibration met the requirements.

The Continuous Calibration File ID BN037641.D met the requirements except for Phenol-d6. Failed surrogate is not associated with DOD, Therefor no further corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.

	
	
	(

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Sample RW5-SP100-20250814 was diluted due to high concentration.

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).”

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

