ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2903

	MATRIX: Water

	METHOD: 8260D

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS  Tuning Specifications

BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY 

ASP CLP, CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.
	
	(
	

	
	The Initial Calibration met the requirements.

The Continuous Calibration File ID VX047511.D met the requirements except for 2-Butanone and Methyl Acetate failing high but no positive hit in associated sample therefore no corrective action taken.
	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for 

MW-1C-081925 [1,2-Dichloroethane-d4 - 130%], MW-1C-081925RE [1,2-Dichloroethane-d4 - 132%, 4-Bromofluorobenzene - 124%], All the failure samples in surrogates was reanalyzed to confirm the failure as per method and reported,

MW-6B-081925MS [1,2-Dichloroethane-d4 - 131%], MW-6B-081925MSD [1,2-Dichloroethane-d4 - 146%, 4-Bromofluorobenzene - 130%, Dibromofluoromethane - 128%, Toluene-d8 - 122%]Surrogate failing for MS-MSD but Original sample passing for that therefore no corrective action taken and 

MW-12B-081925 [4-Bromofluorobenzene - 123%]Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q2903-05MS} with File ID: VX047507.D recoveries met the requirements for all compounds except for 1,3-Dichlorobenzene[83%] due to matrix interference.

The MSD {Q2903-06MSD} with File ID: VX047508.D recoveries met the requirements for all compounds except for 2-Butanone[140%], Bromochloromethane[137%], Chloroform[122%], Methyl Acetate[147%] and Methylene Chloride[119%]due to matrix interference.

The Blank Spike for {VX0825WBS01} with File ID: VX047514.D met requirements for all compounds except for Bromochloromethane[138%]failing high but no positive hit in associated sample therefore no corrective action taken.

The RPD for {Q2903-06MSD} with File ID: VX047508.D met criteria except for Bromomethane[25%]due to difference in results of MS-MSD.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Sample MW-12B-081925 was diluted due to high concentration.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

