ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2921

	MATRIX: Water

	METHOD: 8270E/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The Initial Calibration met the requirements.

The Continuous Calibration File ID BP025561.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde but no positive hit in associated sample therefore no corrective action taken.

The Continuous Calibration File ID BP025579.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde but no positive hit in associated sample therefore no corrective action taken.

The Continuous Calibration File ID BP025596.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde but no positive hit in associated sample therefore no corrective action taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for MW-201-082025MS [Nitrobenzene-d5 - 148%], MW-201-082025MSD [Nitrobenzene-d5 - 149%],as per method one surrogate is allowed to failed, therefore no corrective action was taken.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q2924-03MS} with File ID: BP025602.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[33%], 2,4-Dichlorophenol[157%], 2,4-Dimethylphenol[147%], 2-Methylnaphthalene[-233%], 2-Nitrophenol[168%], 4-Chloro-3-methylphenol[149%], Acetophenone[170%], Hexachlorobutadiene[139%], Hexachloroethane[181%], Isophorone[159%], Naphthalene[-777%] and Nitrobenzene[167%] due to matrix interference.

The MSD {Q2924-04MSD} with File ID: BP025603.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[35%], 2,4-Dichlorophenol[162%], 2,4-Dimethylphenol[149%], 2-Methylnaphthalene[-133%], 2-Nitrophenol[172%], 4-Chloro-3-methylphenol[151%], Acetophenone[172%], Hexachlorobutadiene[142%], Hexachloroethane[184%], Isophorone[162%], Naphthalene[-190%] and Nitrobenzene[169%] due to matrix interference.

The Blank Spike met requirements for all compounds.

The RPD for {Q2924-04MSD} with File ID: BP025603.D met criteria except for 2-Methylnaphthalene[55%], Naphthalene[121%] due to difference in results of MS and MSD.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

