          [image: image2.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
          [image: image1.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2923

	MATRIX: Solid

	METHOD: 8082A/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.


	
	
	(

	2.
	Standard Summary Submitted.


	
	
	(

	3.
	Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD.

The Initial Calibration met the requirements.
The Continuous Calibration File ID PO113428.D met the requirements except for Decachlorobiphenyl is failing in 1st column, however it is passed in 2nd column therefore no corrective action was taken AND Aroclor-1260(Peak-03) is failing in 2nd column however it is passed in 1st column therefore no corrective action was taken.

The Continuous Calibration File ID PP074804.D met the requirements except for Aroclor-1260(Peak-03),Aroclor-1260(Peak-04),Aroclor-1260(Peak-05),Decachlorobiphenyl is failing in 2nd column  however it is passed in 1st column therefore no corrective action was taken.

 ASK  wdcalcomm " "  \* MERGEFORMAT 

	
	
	(

	 4.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	 5.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for K262-DUP3 [Decachlorobiphenyl(1)264%]. As per method one surrogate allowed to fail to meet the criteria per column. No further corrective action was taken.

	
	(
	

	 6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q2914-13MS} with File ID: PO113170.D recoveries met the requirements for all compounds except for [AR1016(1)1010% - AR1016(2)1063%], [AR1260(1)638% - AR1260(2)532%] due to matrix interference.The MS {Q2923-01MS} with File ID: PP074825.D recoveries met the requirements for all compounds except for [AR1016(1)213% - AR1016(2)88%], [AR1260(1)127% - AR1260(2)92%] due to matrix interference.

The MSD {Q2914-13MSD} with File ID: PO113171.D recoveries met the requirements for all compounds except for [AR1016(1)1169% - AR1016(2)1169%], [AR1260(1)691% - AR1260(2)531%]due to matrix interference.

The MSD {Q2923-01MSD} with File ID: PP074826.D recoveries met the requirements for all compounds except for [AR1016(1)206% - AR1016(2)93%], [AR1260(1)126% - AR1260(2)86%] due to matrix interference.

The Blank Spike met requirements for all compounds.

The RPD were met for all analysis.

	
	(
	

	7.
	Retention Time Shift Meet Criteria (if applicable)

Comments: 


	
	
	(

	8.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: 
Sample K262-DUP4 was diluted due to high concentration.

The soil samples  results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date

[image: image1.png][image: image2.png]