ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2924

	MATRIX: Water

	METHOD: 8270E/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Initial Calibration (Method 8270-BF082025.M)  for Hexachlorocyclopentadiene, 2,4-Dinitrophenol  these  Compounds are passing on Linear regression.

The %RSD is greater than 20%  for certain compounds in the Initial Calibration (Method 8270-BF082625.M)  2-Nitroaniline, 2,6-Dinitrotoluene, 3-Nitroaniline, 4-Nitrophenol, 2,4-Dinitrotoluene,4-Nitroaniline, Butylbenzylphthalate, Bis(2-ethylhexyl)phthalate Kept on Linear Regression and Compounds 2-Nitrophenol, 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol, Di-n-octyl phthalate Kept on Quadratic Regression.

The Continuous Calibration File ID BF143596.D met the requirements except for 2-Nitrophenol. Failed high but associated samples does not have hit for this compounds, therefor no further corrective action was taken.

The Continuous Calibration File ID BP025544.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde. Failed high but associated samples does not have hit for this compounds, therefor no further corrective action was taken.

The Continuous Calibration File ID BP025561.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde. Failed high but associated samples does not have hit for this compounds, therefor no further corrective action was taken.

The Continuous Calibration File ID BP025579.D met the requirements except for 2,4-Dinitrophenol and Benzaldehyde. Failed high but associated samples does not have hit for this compounds, therefor no further corrective action was taken.

The Continuous Calibration File ID BP025596.D met the requirements except for 2!4-Dinitrophenol and Benzaldehyde. Failed high but associated samples does not have hit for this compounds, therefor no further corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for MW-201-082025 [Nitrobenzene-d5 - 171%], MW-201-082025DL2 [Phenol-d6 - 9%], MW-201-082025MS [Nitrobenzene-d5 - 148%] and MW-201-082025MSD [Nitrobenzene-d5 - 149%]. As per SOP one Acid and one Base surrogates are allowed to fail, Therefor no further corrective action was taken.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q2924-03MS} with File ID: BP025602.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[33%], 2,4-Dichlorophenol[157%], 2,4-Dimethylphenol[147%], 2-Methylnaphthalene[-233%], 2-Nitrophenol[168%], 4-Chloro-3-methylphenol[149%], Acetophenone[170%], Hexachlorobutadiene[139%], Hexachloroethane[181%], Isophorone[159%], Naphthalene[-777%] and Nitrobenzene[167%]. Recovery failed due to matrix interference, Therefor no further corrective action was taken.

The MSD {Q2924-04MSD} with File ID: BP025603.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[35%], 2,4-Dichlorophenol[162%], 2,4-Dimethylphenol[149%], 2-Methylnaphthalene[-133%], 2-Nitrophenol[172%], 4-Chloro-3-methylphenol[151%], Acetophenone[172%], Hexachlorobutadiene[142%], Hexachloroethane[184%], Isophorone[162%], Naphthalene[-190%] and Nitrobenzene[169%]. Recovery failed due to matrix interference, Therefor no further corrective action was taken.

The Blank Spike met requirements for all compounds.

The RPD for {Q2924-04MSD} with File ID: BP025603.D met criteria except for 2-Methylnaphthalene[55%], Naphthalene[121%]. RPD failed due to result diffrance between MS and MSD results , Therefor no further corrective action was taken.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples MW-201-082025, MW-201-082025DL, MW-2B-082025 and MW-2B-082025DL were diluted due to high concentrations.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

