ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q2924

	MATRIX: Water

	METHOD: 8260D

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS  Tuning Specifications

BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY 

ASP CLP, CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.
	
	(
	

	
	The %RSD is greater than 20% in the Initial Calibration method (82V082125W.M) for Bromochloromethane, 2-Hexanone this compounds are passing on Quadratic regression.

The Continuous Calibration File ID VV038982.D met the requirements except for Methyl Acetate is failing high but no positive hit in associate sample therefore no corrective action taken while  Styrene is failing marginally high therefore no corrective action taken.

The Continuous Calibration File ID VV039009.D met the requirements except for Bromomethane and Chloromethane are failing low, while 2-Butanone,Acetone, Methyl Acetate and Methyl tert-butyl Ether failing high but only dilution sample analyzed under this CCAL and dilution not required for mentioned analyte therefore no corrective action taken.
	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q2924-03MS} with File ID: VV039005.D recoveries met the requirements for all compounds except for 1,1,2,2-Tetrachloroethane[58%], 1,2-Dibromoethane[80%], 1,2-Dichloroethane[360%], 1,2-Dichloropropane[82%], 2-Hexanone[64%], Benzene[-200%], cis-1,2-Dichloroethene[40%], Ethyl Benzene[0%], Isopropylbenzene[70%], m/p-Xylenes[-100%], o-Xylene[-600%], Styrene[-200%], Toluene[-200%] and trans-1,2-Dichloroethene[20%] due to matrix interference.

The MSD {Q2924-04MSD} with File ID: VV039006.D recoveries met the requirements for all compounds except for 1,2-Dichloroethane[360%], 2-Hexanone[68%], Benzene[-200%], cis-1,2-Dichloroethene[60%], Ethyl Benzene[0%], m/p-Xylenes[-100%], o-Xylene[0%], Styrene[0%] and Toluene[-400%] due to matrix interference.

The Blank Spike Duplicate for {VV0822WBS01} with File ID: VV039011.D met requirements for all compounds except for Bromochloromethane[126%], Methyl Acetate[135%], Methyl tert-butyl Ether[118%] and trans-1,2-Dichloroethene[109%] are failing high but no positive hit in associate sample therefore no corrective action taken.

The RPD for {Q2924-04MSD} with File ID: VV039006.D met criteria except for 1,1,2,2-Tetrachloroethane[50%], 1,2-Dibromo-3-Chloropropane[29%], Acetone[21%], Bromoform[32%], Bromomethane[22%], Chloromethane[41%], cis-1,2-Dichloroethene[40%], Isopropylbenzene[42%], Methyl Acetate[26%], Methylene Chloride[21%], Toluene[67%], trans-1,2-Dichloroethene[120%] and Vinyl chloride[37%] due to difference in results of MS and MSD.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples MW-201-082025, MW-201-082025DL, MW-2B-082025 and MW-2B-082025DL were diluted due to high concentrations.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

