[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Zuccaro Inc.
Project Name: 3 Coal Ave., Morristown, NJ
Project # N/A
Order ID # Q3032
Test Name: SVOC-TCL BNA -20
A. Number of Samples and Date of Receipt:

6 Solid samples were received on 09/05/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

Cyanide, Mercury, Metals ICP-TAL, PCB, Pesticide-TCL, SVOC-TCL BNA -20, TCL+30/TAL, TCLP BNA, TCLP Extraction, TCLP Herbicide, TCLP ICP Metals, TCLP Mercury, TCLP Pesticide, TCLP VOA, TCLP ZHE Extraction, TCLP-FULL and VOC-TCLVOA-10. This data package contains results for SVOC-TCL BNA -20.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_G using GC Column ZB-Semi Volatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGAThe analysis of SVOC-TCL BNA -20 was based on method 8270E and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries met the acceptable criteria.
The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS {Q3032-01MS} with File ID: BG064310.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[65%], 3,3-Dichlorobenzidine[42%], 3-Nitroaniline[50%] and 4-Chloroaniline[29%], these compounds did not meet the NJDKQP criteria but met the in-house criteria  due to matrix interference.
The MSD {Q3032-01MSD} with File ID: BG064311.D recoveries met the acceptable requirements except for 2,2-oxybis(1-Chloropropane)[65%], 3,3-Dichlorobenzidine[43%], 3-Nitroaniline[55%] and 4-Chloroaniline[32%], these compounds did not meet the NJDKQP criteria but met the in-house criteria  due to matrix interference.
The RPD met criteria .
The Blank Spike for {PB169593BS} with File ID: BG064331.D met requirements for all samples except for 2,2-oxybis(1-Chloropropane)[59%], 3,3-Dichlorobenzidine[65%] and 4-Chloroaniline[54%], these compounds did not meet the NJDKQP criteria but met the in-house criteria .

The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Method 8270-BG082825.M for 2,4-Dinitrophenol & it is passing on Linear regression.

The Continuous Calibration File ID BG064301.D met the requirements except for 2,2-oxybis(1-Chloropropane), marginally biased low and 4,6-Dinitro-2-methylphenol,Hexachlorobutadiene,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol failing biased high, The associate samples have no positive hit for these compounds therefore no corrective action was taken. 

The Continuous Calibration File ID BG064329.D met the requirements except for 1,2,4,5-Tetrachlorobenzene, 2,3,4,6-Tetrachlorophenol, 2,4,5-Trichlorophenol, 2,4-Dichlorophenol, 2,4-Dinitrophenol, 2-Nitrophenol, 4,6-Dinitro-2-methylphenol, 4-Bromophenyl-phenylether, 4-Nitroaniline, Atrazine, Hexachlorobutadiene, Hexachlorocyclopentadiene, Pentachlorophenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken.and 2,4,6-Tribromophenol, biased high therefore no corrective action was taken.
The Tuning criteria met requirements. 
Samples SOIL-PILE-2 was diluted due to viscous and concentrated matrix.
E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

The soil samples results are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


