ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3032

	MATRIX: Solid

	METHOD: 8270E/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BG082825.M for 2,4-Dinitrophenol & it is passing on Linear regression.

The Continuous Calibration File ID BG064301.D met the requirements except for 2,2-oxybis(1-Chloropropane), marginally biased low and 4,6-Dinitro-2-methylphenol,Hexachlorobutadiene,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol failing biased high, The associate samples have no positive hit for these compounds therefore no corrective action was taken. 

The Continuous Calibration File ID BG064329.D met the requirements except for 1,2,4,5-Tetrachlorobenzene, 2,3,4,6-Tetrachlorophenol,2,4,5-Trichlorophenol, 2,4-Dichlorophenol, 2,4-Dinitrophenol,2-Nitrophenol, 4,6-Dinitro-2-methylphenol, 4-Bromophenyl-phenylether, 4-Nitroaniline,Atrazine, Hexachlorobutadiene, Hexachlorocyclopentadiene, Pentachlorophenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. . and 2,4,6-Tribromophenol, biased high therefore no corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3032-01MS} with File ID: BG064310.D recoveries met the requirements for all compounds except for 2,2-oxybis(1-Chloropropane)[65%], 3,3-Dichlorobenzidine[42%], 3-Nitroaniline[50%] and 4-Chloroaniline[29%], these compounds did not meet the NJDKQP criteria but met the in-house criteria  due to matrix interference.

The MSD {Q3032-01MSD} with File ID: BG064311.D recoveries met the acceptable requirements except for 2,2-oxybis(1-Chloropropane)[65%], 3,3-Dichlorobenzidine[43%], 3-Nitroaniline[55%] and 4-Chloroaniline[32%], these compounds did not meet the NJDKQP criteria but met the in-house criteria  due to matrix interference.

The Blank Spike for {PB169593BS} with File ID: BG064331.D met requirements for all samples except for 2,2-oxybis(1-Chloropropane)[59%], 3,3-Dichlorobenzidine[65%] and 4-Chloroaniline[54%], these compounds did not meet the NJDKQP criteria but met the in-house criteria .


	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples SOIL-PILE-2 was diluted due to viscous and concentrated matrix.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.The not QT review data is reported in the Miscellaneous.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
_______________________________________



__________________________

QA REVIEW







Date

