[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Apex Companies, LLC.
Project Name: Emerson Vapor Intrusion REG006-0309010-25011690
Project # N/A
Order ID # Q3037
Test Name: VOCMS Group2
A. Number of Samples and Date of Receipt:

1 Air sample was received on 09/05/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group2. This data package contains results for VOCMS Group2.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_L were done using GC column RTX-1, which is 60 meters, 0.32 mm id, 1.0 um df, Restek Cat. #10157. The Trap was supplied by Entech, glass bead and Tenax , Entech 7100A Preconcentrator.The analysis of VOCMS Group2 was based on method TO-15.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
The not QT review data is reported in the Miscellaneous.

The Manual Integrations are performed for the followings:
	Manual Integration Report

	Sequence
	VL081325
	Instrument
	MSVOA_l
	
	
	
	


	Sample ID
	File ID
	Parameter
	Review By
	Review On
	Supervised By
	Supervised On
	Reason


	VSTDICCC010
	VL042842.D
	m/p-Xylene
	SAM
	8/14/2025 7:48:06 AM
	MMDadoda
	8/18/2025 8:42:07 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	1,1,2-Trichloroethane
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	1,4-Dioxane
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	cis-1,3-Dichloropropene
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	Ethanol
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	m/p-Xylene
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042843.D
	Methyl Methacrylate
	SAM
	8/14/2025 7:48:12 AM
	MMDadoda
	8/18/2025 8:42:11 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	1,2-Dichloropropane
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	1,4-Dioxane
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	2,2,4-Trimethylpentane
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	Ethanol
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	Heptane
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	m/p-Xylene
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	Methyl Methacrylate
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	t-1,3-Dichloropropene
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042844.D
	trans-1,2-Dichloroethene
	SAM
	8/14/2025 7:48:18 AM
	MMDadoda
	8/18/2025 8:42:13 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	1,1,2-Trichloroethane
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	1,4-Dioxane
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	2,2,4-Trimethylpentane
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	4-Methyl-2-Pentanone
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	cis-1,3-Dichloropropene
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	m/p-Xylene
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	Methyl Methacrylate
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	t-1,3-Dichloropropene
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042845.D
	Toluene
	SAM
	8/14/2025 7:49:25 AM
	MMDadoda
	8/18/2025 8:42:16 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042846.D
	1,1,1-Trichloroethane
	SAM
	8/14/2025 7:48:23 AM
	MMDadoda
	8/18/2025 8:42:18 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042846.D
	1,2-Dibromoethane
	SAM
	8/14/2025 7:48:23 AM
	MMDadoda
	8/18/2025 8:42:18 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042846.D
	Naphthalene
	SAM
	8/14/2025 7:48:23 AM
	MMDadoda
	8/18/2025 8:42:18 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042846.D
	Tetrachloroethene
	SAM
	8/14/2025 7:48:23 AM
	MMDadoda
	8/18/2025 8:42:18 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042846.D
	Trichloroethene
	SAM
	8/14/2025 7:48:23 AM
	MMDadoda
	8/18/2025 8:42:18 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042847.D
	1,1,1-Trichloroethane
	SAM
	8/14/2025 7:49:29 AM
	MMDadoda
	8/18/2025 8:42:20 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042847.D
	1,1,2,2-Tetrachloroethane
	SAM
	8/14/2025 7:49:29 AM
	MMDadoda
	8/18/2025 8:42:20 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042847.D
	Carbon Tetrachloride
	SAM
	8/14/2025 7:49:29 AM
	MMDadoda
	8/18/2025 8:42:20 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042847.D
	Tetrachloroethene
	SAM
	8/14/2025 7:49:29 AM
	MMDadoda
	8/18/2025 8:42:20 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042847.D
	Trichloroethene
	SAM
	8/14/2025 7:49:29 AM
	MMDadoda
	8/18/2025 8:42:20 AM
	Peak Integrated by Software incorrectly

	VSTDICC015
	VL042848.D
	m/p-Xylene
	SAM
	8/14/2025 7:49:35 AM
	MMDadoda
	8/18/2025 8:42:35 AM
	Peak Integrated by Software incorrectly

	VSTDICV010
	VL042849.D
	cis-1,3-Dichloropropene
	SAM
	8/14/2025 7:48:28 AM
	MMDadoda
	8/18/2025 8:42:37 AM
	Peak Integrated by Software incorrectly

	VSTDICV010
	VL042849.D
	m/p-Xylene
	SAM
	8/14/2025 7:48:28 AM
	MMDadoda
	8/18/2025 8:42:37 AM
	Peak Integrated by Software incorrectly


	Sequence
	VL090825
	Instrument
	MSVOA_l
	
	
	
	


	Sample ID
	File ID
	Parameter
	Review By
	Review On
	Supervised By
	Supervised On
	Reason


	VSTDCCC010
	VL042900.D
	m/p-Xylene
	sam
	9/9/2025 12:26:34 PM
	MMDadoda
	9/11/2025 5:28:20 PM
	Peak Integrated by Software incorrectly

	VL0908ABS01
	VL042902.D
	m/p-Xylene
	sam
	9/9/2025 12:26:38 PM
	MMDadoda
	9/11/2025 5:24:29 PM
	Peak Integrated by Software incorrectly

	VL0908ABS01
	VL042902.D
	Propene
	sam
	9/9/2025 12:26:38 PM
	MMDadoda
	9/11/2025 5:24:29 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	4-Methyl-2-Pentanone
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Benzene
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Carbon Tetrachloride
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	m/p-Xylene
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Methyl Methacrylate
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	n-propylbenzene
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Propene
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Trichlorofluoromethane
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01
	VL042904.D
	Vinyl Acetate
	sam
	9/9/2025 12:26:43 PM
	MMDadoda
	9/11/2025 5:28:18 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Benzene
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	m/p-Xylene
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Methyl Methacrylate
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	n-propylbenzene
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Propene
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Tetrahydrofuran
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Toluene
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Trichlorofluoromethane
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly

	Q3037-01DUP
	VL042905.D
	Vinyl Acetate
	sam
	9/9/2025 12:31:10 PM
	MMDadoda
	9/11/2025 5:28:16 PM
	Peak Integrated by Software incorrectly


F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


