ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GENERAL CHEMISTRY CONFORMANCE/NON-CONFORMANCE SUMMARY
CHEMTECH 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GENERAL CHEMISTRY CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)

	ORDER ID: Q3041

	MATRIX: Water

	METHOD: 1664A,SM2540 D,SM4500-NH3,SM5210 B

	


	
	
	NA
	NO
	YES

	   1.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	   2.
	Matrix Spike Duplicate Recoveries Met Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike met requirements for all compounds.

	
	
	(

	   3.
	Sample Duplicate Analysis Met QC Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	   8.
	Digestion Holding Time Met 






If not met, list number of days exceeded for each sample:

The Holding Times were met for all samples except for 001 Willets Pt Blvd (Sep) of BOD5 and for 002 35th Ave(sep) of BOD5 as samples were receive out of holding time.

	
	(
	


ADDITIONAL COMMENTS: The temperature of the samples at the time of receipt was 20.1°C. Lab notified this issue to the client. See the communication in shipping Document section. 

As per method 1664A, MS/MSD is required to be performed with the sample analysis. However, Lab did not receive sufficient volume to perform the MS/MSD for Q3041 therefore Lab reported MS-MSD from Q3057.
_______________________________________



__________________________

QA REVIEW







Date

