[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

G Environmental
Project Name: Buffington
Project # N/A
Order ID # Q3058
Test Name: VOCMS Group1,VOC-TCLVOA-10
A. Number of Samples and Date of Receipt:

2 Solid samples were received on 09/09/2025.

1 Water sample was received on 09/09/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for VOCMS Group1(8260-Low), VOC-TCLVOA-10(8260D).

C. Analytical Techniques:

VOCMS Group1 : The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group1 was based on method 8260-Low.

VOC-TCLVOA-10 : The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-TCLVOA-10 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis. 

The Surrogate recoveries were met for all analysis except following 

VOC-TCLVOA-10 : RPXY4 [1,2-Dichloroethane-d4 - 136%] this compound did not meet the NJDKQP criteria but met the in-house criteria.

The Internal Standards Areas were met for all analysis. 

The Retention Times were met for all analysis. 

The RPD were met for all analysis. 

The Blank Spike met requirements for all compounds.  
The Blank Spike Duplicate met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination. 

The Initial Calibration met the requirements except following 

VOCMS Group1 : The %RSD is greater than 20% in the Initial Calibration method (82N082125W.M) for Methylene Chloride passing on Linear regression.

VOC-TCLVOA-10 : The %RSD is greater than 20% in the Initial Calibration method (82N082125W.M) for Methylene Chloride passing on Linear regression.

The Continuous Calibration met the requirements. 

The Tuning criteria met requirements. 

VOC-TCLVOA-10 : Sample RPXY5 was diluted due to high concentration.
E. Additional Comments:
Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples results are based on a dry weight basis.

The Sample # RPXY4 have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


