[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

G Environmental
Project Name: Blanchard
Project # N/A
Order ID # Q3083
Test Name: SVOC-SIM
A. Number of Samples and Date of Receipt:

1 Water sample was received on 09/11/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-SIM. This data package contains results for SVOC-SIM.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_N using GC Column ZB-Semi Volatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-SIM was based on method 8270-Modified and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for MW3SDL [Terphenyl-d14 - 162%], this compound did not meet the NJDKQP criteria but met the in-house criteria, as per method one acid and one base surrogate allow to fail therefore no corrective action taken, and MW3S [Terphenyl-d14 - 199%], this compound did not meet the NJDKQP criteria and in-house criteria, 
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the requirements for all compounds.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

Sample MW3S was diluted due to high concentration.
E. Additional Comments:

The Sample MW3S, and MW3SDL, have the concentration of target compound below method detection limits; therefore it is not reported as Hit in Form1.
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


