ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3097

	MATRIX: Water

	METHOD: 8260D

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS  Tuning Specifications

BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY 

ASP CLP, CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.
	
	(
	

	
	The Initial Calibration met the requirements.
The Continuous Calibration File ID VX047660.D met the requirements except for Bromomethane and Dichlorodifluoromethane are failing high but no positive hit in associate sample therefore no corrective action taken.

The Continuous Calibration File ID VX047689.D met the requirements except for Bromomethane,Carbon Disulfide,Chloromethane,Dichlorodifluoromethane and Vinyl Chloride are failing high but no positive hit in associate sample therefore no corrective action taken.

The Continuous Calibration File ID VX047719.D met the requirements except for Bromomethane,Dichlorodifluoromethane and Vinyl Chloride are failing high and associate sample having hit of Dichlorodifluoromethane  but below CRQL therefore no corrective action taken.

The Continuous Calibration File ID VX047746.D met the requirements except for Dichlorodifluoromethane is failing high but no positive hit in associate sample therefore no corrective action taken.

The Continuous Calibration File ID VX047769.D met the requirements except for Methyl Acetate  is failing high but no positive hit in associate sample therefore no corrective action taken.

END CCAL was not analyzed in sequence VX092425 due to power error in laboratory during the night and there are no more vials for re-analysis therefore all samples reported from this sequence without END CCAL.
	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for 

RE103D2-20250911 [1,2-Dichloroethane-d4 - 122%], 

RE108D2-20250911DL [1,2-Dichloroethane-d4 - 119%, 4-Bromofluorobenzene - 118% and Toluene-d8 - 119%], Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported while for 

RE103D2-20250911MS [Toluene-d8 - 113%], surrogate recoveries passing in parent sample and MSD therefore no corrective action taken.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3097-02MS} with File ID: VX047715.D recoveries met the requirements for all compounds except for Bromochloromethane[124%], Trichloroethene[200%] due to matrix interference.

The MSD {Q3097-03MSD} with File ID: VX047716.D recoveries met the requirements for all compounds except for Trichloroethene[140%] due to matrix interference. 
The sample # RE103D2-20250911MS and RE103D2-20250911MSD is failing for Bromochloromethane and Trichloroethene therefore the original sample(RE103D2-20250911) is reported with M flag for this compounds. 

The RPD for {Q3097-03MSD} with File ID: VX047716.D met criteria except for Trichloroethene[35%] due to difference in results of MS and MSD.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples RE103D2-20250911, RE108D2-20250911, RE125D2-20250911, RE105D2-20250911, RE120D1-20250912, DUP06-20250912, RE120D2-20250912, RE134D4-20250912 and RE134D3-20250912 were diluted due to high concentrations.

The Sample #RE125D1-20250911, RE105D1-20250911, DUP04-20250911, RE120D1-20250912, RE134D1-20250912, RE134D2-20250912 have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1.

The not QT review data is reported in the Miscellaneous.

_______________________________________



__________________________

QA REVIEW







Date

