          [image: image2.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
          [image: image1.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3099

	MATRIX: Solid

	METHOD: 8082A/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.


	
	
	(

	2.
	Standard Summary Submitted.


	
	
	(

	3.
	Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD.

The Initial Calibration met the requirements.
The Continuous Calibration File ID PO113755.D met the requirements except for Aroclor-1260(Peak-05) is failing in 1st column, however it is passed in 2nd column therefore no corrective action was taken.

The Continuous Calibration File ID PO113770.D met the requirements except for Aroclor-1260(Peak-04) is failing in 1st column, however it is passed in 2nd column therefore no corrective action was taken.

 ASK  wdcalcomm " "  \* MERGEFORMAT 

	
	
	(

	 4.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	 5.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	 6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike met requirements for all compounds.


	
	
	(

	7.
	Retention Time Shift Meet Criteria (if applicable)

Comments: 


	
	
	(

	8.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: 
Samples 6A6B6C was diluted due to bad matrix.

The temperature were not listed in the COC. 

No MSMSD performed as samples are caulk matrix.

Less volume was taken at the time of extraction due to caulk matrix of the samples.

The soil samples  results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date

[image: image1.png][image: image2.png]