[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

RTP Environmental
Project Name: CHADWICK SQUARE
Project # N/A
Order ID # Q3132
Test Name: TO-15
A. Number of Samples and Date of Receipt:

2 Air samples were received on 09/17/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

TO-15. This data package contains results for TO-15.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_L were done using GC column RTX-1, which is 60 meters, 0.32 mm id, 1.0 um df, Restek Cat. #10157. The Trap was supplied by Entech, glass bead and Tenax , Entech 7100A Preconcentrator.The analysis of TO-15 was based on method TO-15.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD for {Q3130-02DUP} with File ID: VL042970.D met criteria except for 1,2-Dichloroethane[30.8%], Tetrachloroethene[200%] due to difference in results of original and DUP.
The Blank Spike met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 30% in the Initial Calibration method (VL092225AIR.M) for Methylene Chloride passing on Linear Regression.

The Tuning criteria met requirements.          

Sample 151-1A-1 was diluted due to high concentration.
E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

Alliance Technical Group, LLC - Newark holds certification for the analytes " Bromodichloromethane, Chloroethane and Heptane". Currently, we are not eligible to report them, as we have not successfully completed the required Proficiency Testing (PT) study. We are actively engaged in the PT study process to gain reporting eligibility.
The Sample #151-1A-1 and 151-1A-1DL have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1.
The Manual Integrations are performed for the followings.
	Manual Integration Report

	Sequence
	VL092225
	Instrument
	MSVOA_l
	
	
	
	


	Sample ID
	File ID
	Parameter
	Review By
	Review On
	Supervised By
	Supervised On
	Reason

	VSTDICCC010
	VL042950.D
	m/p-Xylene
	sam
	9/23/2025 10:01:03 AM
	MMDadoda
	9/24/2025 2:16:17 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042951.D
	1,1,2-Trichloroethane
	sam
	9/23/2025 10:01:08 AM
	MMDadoda
	9/24/2025 2:16:19 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042951.D
	1,2-Dichloropropane
	sam
	9/23/2025 10:01:08 AM
	MMDadoda
	9/24/2025 2:16:19 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042951.D
	1,4-Dioxane
	sam
	9/23/2025 10:01:08 AM
	MMDadoda
	9/24/2025 2:16:19 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042951.D
	4-Methyl-2-Pentanone
	sam
	9/23/2025 10:01:08 AM
	MMDadoda
	9/24/2025 2:16:19 AM
	Peak Integrated by Software incorrectly

	VSTDICC002
	VL042951.D
	m/p-Xylene
	sam
	9/23/2025 10:01:08 AM
	MMDadoda
	9/24/2025 2:16:19 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	1,2-Dichloropropane
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	1,4-Dioxane
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	Bromodichloromethane
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	cis-1,3-Dichloropropene
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	Heptane
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	m/p-Xylene
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC001
	VL042952.D
	t-1,3-Dichloropropene
	sam
	9/23/2025 10:01:13 AM
	MMDadoda
	9/24/2025 2:16:21 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	1,1,2-Trichloroethane
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	1,4-Dioxane
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	2,2,4-Trimethylpentane
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	Benzyl Chloride
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	cis-1,3-Dichloropropene
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	Ethanol
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	Heptane
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	m/p-Xylene
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	Methyl Methacrylate
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	t-1,3-Dichloropropene
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	Tetrachloroethene
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.5
	VL042953.D
	trans-1,2-Dichloroethene
	sam
	9/23/2025 10:02:11 AM
	MMDadoda
	9/24/2025 2:16:23 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042954.D
	1,2-Dibromoethane
	sam
	9/23/2025 10:02:56 AM
	MMDadoda
	9/24/2025 2:16:25 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042954.D
	Carbon Tetrachloride
	sam
	9/23/2025 10:02:56 AM
	MMDadoda
	9/24/2025 2:16:25 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042954.D
	Naphthalene
	sam
	9/23/2025 10:02:56 AM
	MMDadoda
	9/24/2025 2:16:25 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042954.D
	Tetrachloroethene
	sam
	9/23/2025 10:02:56 AM
	MMDadoda
	9/24/2025 2:16:25 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.1
	VL042954.D
	Trichloroethene
	sam
	9/23/2025 10:02:56 AM
	MMDadoda
	9/24/2025 2:16:25 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042955.D
	1,1,1-Trichloroethane
	sam
	9/23/2025 10:01:57 AM
	MMDadoda
	9/24/2025 2:16:26 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042955.D
	1,1,2,2-Tetrachloroethane
	sam
	9/23/2025 10:01:57 AM
	MMDadoda
	9/24/2025 2:16:26 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042955.D
	Carbon Tetrachloride
	sam
	9/23/2025 10:01:57 AM
	MMDadoda
	9/24/2025 2:16:26 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042955.D
	Tetrachloroethene
	sam
	9/23/2025 10:01:57 AM
	MMDadoda
	9/24/2025 2:16:26 AM
	Peak Integrated by Software incorrectly

	VSTDICC0.03
	VL042955.D
	Trichloroethene
	sam
	9/23/2025 10:01:57 AM
	MMDadoda
	9/24/2025 2:16:26 AM
	Peak Integrated by Software incorrectly

	VSTDICC015
	VL042956.D
	m/p-Xylene
	sam
	9/23/2025 10:01:19 AM
	MMDadoda
	9/24/2025 2:16:27 AM
	Peak Integrated by Software incorrectly

	VSTDICV010
	VL042957.D
	m/p-Xylene
	sam
	9/23/2025 10:01:52 AM
	MMDadoda
	9/24/2025 2:16:29 AM
	Peak Integrated by Software incorrectly

	VL0922ABS01
	VL042959.D
	m/p-Xylene
	sam
	9/23/2025 10:01:29 AM
	MMDadoda
	9/24/2025 2:16:44 AM
	Peak Integrated by Software incorrectly

	Q3130-02DUP
	VL042970.D
	2,2,4-Trimethylpentane
	sam
	9/23/2025 10:03:16 AM
	MMDadoda
	9/24/2025 2:17:13 AM
	Peak Integrated by Software incorrectly

	Q3130-02DUP
	VL042970.D
	Carbon Tetrachloride
	sam
	9/23/2025 10:03:16 AM
	MMDadoda
	9/24/2025 2:17:13 AM
	Peak Integrated by Software incorrectly

	Q3130-02DUP
	VL042970.D
	m/p-Xylene
	sam
	9/23/2025 10:03:16 AM
	MMDadoda
	9/24/2025 2:17:13 AM
	Peak Integrated by Software incorrectly

	Q3130-02DUP
	VL042970.D
	Propene
	sam
	9/23/2025 10:03:16 AM
	MMDadoda
	9/24/2025 2:17:13 AM
	Peak Integrated by Software incorrectly

	Q3130-02DUP
	VL042970.D
	Tetrachloroethene
	sam
	9/23/2025 10:03:16 AM
	MMDadoda
	9/24/2025 2:17:13 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	2,2,4-Trimethylpentane
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Acetone
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Carbon Tetrachloride
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Chlorodifluoromethane
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Cyclohexane
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Propene
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Styrene
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	tert-Butyl alcohol
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Tetrachloroethene
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01
	VL042978.D
	Toluene
	MMDadoda
	9/24/2025 2:17:55 AM
	SAM
	9/24/2025 2:25:43 AM
	Peak Integrated by Software incorrectly

	Q3132-01DL
	VL042979.D
	Acetone
	MMDadoda
	9/24/2025 2:17:57 AM
	SAM
	9/24/2025 2:25:44 AM
	Peak Integrated by Software incorrectly

	Q3132-01DL
	VL042979.D
	m/p-Xylene
	MMDadoda
	9/24/2025 2:17:57 AM
	SAM
	9/24/2025 2:25:44 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	2,2,4-Trimethylpentane
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	4-Methyl-2-Pentanone
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Acetone
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Carbon Tetrachloride
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Chlorodifluoromethane
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Cyclohexane
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Heptane
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly

	Q3132-02
	VL042980.D
	Propene
	MMDadoda
	9/24/2025 2:17:58 AM
	SAM
	9/24/2025 2:25:46 AM
	Peak Integrated by Software incorrectly


F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


