[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Scalamandre – Tully JV
Project Name: NYC DOT Harper Street Yard North
Project # N/A
Order ID # Q3150
Test Name: VOCMS Group1
A. Number of Samples and Date of Receipt:

4 Solid samples were received on 09/19/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group1. This data package contains results for VOCMS Group1.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_W were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group1 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis except for MER Rd Con Ed, MER Rd Con EdRE, Mid Site Grid 5-7 and Mid Site Grid 5-7RE the failure samples were reanalyzed to confirm the failure and reported.
The Retention Times were met for all analysis.
The RPD for {VW0919SBSD01} with File ID: VW032207.D met criteria except for Trichlorofluoromethane[24%]. Due to difference in BS and BSD concentrations.
The Blank Spike for {VW0919SBS01} with File ID: VW032206.D met requirements for all compounds except for 1,1,2-Trichlorotrifluoroethane[125%], 1,1-Dichloroethene[123%].Failing high but associated samples have no positive hit for these compounds therefore no corrective action was taken.

The Blank Spike Duplicate for {VW0919SBSD01} with File ID: VW032207.D met requirements for all compounds except for 1,1-Dichloroethene[127%] and Chloroethane[142%]. Failing high but associated samples have no positive hit for these compounds therefore no corrective action was taken.

The Blank Spike for {VW0922SBS01} with File ID: VW032225.D met requirements for all compounds except for Chloroethane[132%]. Failing high but associated samples have no positive hit for this compound therefore no corrective action was taken.

The Blank Spike for {VW1002SBS01} with File ID: VW032324.D met requirements for all compounds except for Methylene Chloride[134%]. failing high but associated sample having hit for Methylene Chloride  below CRQL therefore no corrective action was taken.

The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Initial Calibration method (82W082625S.M) for Methylene chloride, this compound is passing on Quadratic regression.

The %RSD is greater than 20% in the Initial Calibration method (82W092525S.M) for Methylene chloride, this compound is passing on Linear regression.

The %RSD is greater than 20% in the Initial Calibration method (82W100125S.M) for Methylene chloride, this compound is passing on Linear regression.

The Continuous Calibration File ID VW032204.D met the requirements except for Methyl Acetate. Failing marginally therefore no corrective action was taken.
The Tuning criteria met requirements.          

E. Additional Comments:
Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

The temperature of the samples at the time of receipt was 18.3°C
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


