ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3150

	MATRIX: Solid

	METHOD: 8260D

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS  Tuning Specifications

BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY 

ASP CLP, CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.
	
	
	(

	
	The %RSD is greater than 20% in the Initial Calibration method (82W082625S.M) for Methylene chloride, this compound is passing on Quadratic regression.

The %RSD is greater than 20% in the Initial Calibration method (82W092525S.M) for Methylene chloride, this compound is passing on Linear regression.

The %RSD is greater than 20% in the Initial Calibration method (82W100125S.M) for Methylene chloride, this compound is passing on Linear regression.

The Continuous Calibration File ID VW032204.D met the requirements except for Methyl Acetate. Failing marginally therefore no corrective actione was taken.
	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike for {VW0919SBS01} with File ID: VW032206.D met requirements for all compounds except for 1,1,2-Trichlorotrifluoroethane[125%], 1,1-Dichloroethene[123%].Failing high but associated samples have no positive hit for these compounds therefore no corrective action was taken.

The Blank Spike Duplicate for {VW0919SBSD01} with File ID: VW032207.D met requirements for all compounds except for 1,1-Dichloroethene[127%] and Chloroethane[142%]. Failing high but associated samples have no positive hit for these compounds therefore no corrective action was taken.

The Blank Spike for {VW0922SBS01} with File ID: VW032225.D met requirements for all compounds except for Chloroethane[132%]. Failing high but associated samples have no positive hit for this compound therefore no corrective action was taken.

The Blank Spike for {VW1002SBS01} with File ID: VW032324.D met requirements for all compounds except for Methylene Chloride[134%]. failing high but associated sample having hit for Methylene Chloride  below CRQL therefore no corrective action was taken.

The Blank Spike Duplicate for {VW0922SBSD01} with File ID: VW032226.D met requirements for all compounds except for 1,1,2-Trichlorotrifluoroethane[136%], 1,1-Dichloroethene[133%], Chloroethane[138%] and Trichlorofluoromethane[138%].

The RPD for {VW0919SBSD01} with File ID: VW032207.D met criteria except for Trichlorofluoromethane[24%]. Due to difference in BS and BSD concentrations.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: The Internal Standards Areas were met for all analysis except for MER Rd Con Ed, MER Rd Con EdRE, Mid Site Grid 5-7 and Mid Site Grid 5-7RE the failure samples were reanalyzed to confirm the failure and reported.

	
	(
	

	10.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

The temperature of the samples at the time of receipt was 18.3°C
_______________________________________



__________________________

QA REVIEW







Date

