[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Always Available Construction LLC.
Project Name: Yard Soil Cleanup
Project # N/A
Order ID # Q3157
Test Name: SVOC-TCL BNA -20,TCLP BNA
A. Number of Samples and Date of Receipt:

2 Solid samples were received on 09/19/2025.

1 Water sample was received on 09/19/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for SVOC-TCL BNA -20(8270E), TCLP BNA(8270E).

C. Analytical Techniques:

TCLP BNA : The samples were analyzed on instrument BNA_G using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of TCLP BNA was based on method 8270E and extraction was done based on method 3510 and TCLP extraction method was 1311.

SVOC-TCL BNA -20 : The samples were analyzed on instrument BNA_P using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-TCL BNA -20 was based on method 8270E and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis. 
The Surrogate recoveries were met for all analysis except following 

TCLP BNA : VNJ-238MS [2,4,6-Tribromophenol - 147%], VNJ-238MSD [2,4,6-Tribromophenol - 151%], 20-DEAD-1 [2,4 and6-Tribromophenol - 156%], PB169754TB [2,4 and6-Tribromophenol - 138%].  AS per method one Acid surrogate is allowed to fail, Therefore no corrective action required.

The Internal Standards Areas were met for all analysis. 
The Retention Times were met for all analysis. 

The MS recoveries met the requirements for all compounds except following 

TCLP BNA : The MS {Q3133-02MS} with File ID: BG064423.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[124%], 2,4,6-Trichlorophenol[116%] and Hexachlorobenzene[124%] due to matrix interference.

The MSD recoveries met the requirements for all compounds except following 

TCLP BNA : The MSD {Q3133-02MSD} with File ID: BG064424.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[124%], 2,4,6-Trichlorophenol[124%] and Hexachlorobenzene[124%] due to matrix interference.

The RPD were met for all analysis. 
The Blank Spike met requirements for all compounds except following 
TCLP BNA : The Blank Spike for {PB169783BS} with File ID: BG064420.D met requirements for all compounds except for Hexachlorobutadiene[103%]. But associaed samples have not positive hit for this compound therefore no corrective action was taken.

The Blank analysis did not indicate the presence of lab contamination. 

The Initial Calibration met the requirements except following 

SVOC-TCL BNA -20 : The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP091225.M)  for 2,4-Dinitrophenol & 4-Nitrophenol.  These Compounds are passing on Linear regression.

The Continuous Calibration met the requirements except following 

TCLP BNA : The Continuous Calibration File ID BG064418.D met the requirements except for 1,4-Dichlorobenzene,2,4,5-Trichlorophenol,2,4,6-Trichlorophenol,2,4-Dinitrotoluene, Hexachlorobenzene, Hexachlorobutadiene,Pentachlorophenol,2,4,6-Tribromophenol,2-Fluorophenol,Nitrobenzene-d5 and Terphenyl-d14. But associaed samples have not positive hit for these compounds therefore no corrective action was taken.

SVOC-TCL BNA -20 : The Continuous Calibration File ID BP025824.D met the requirements except for 4-Nitrophenol,Benzaldehyde,Benzo(g,h,i)perylene and Caprolactam. But associated samples have not positive hit for these compounds therefore no corrective acrion was taken.

The Tuning criteria met requirements. 

E. Additional Comments:
The temperature of the samples at the time of receipt was 26.1°C.

The soil samples  results  are based on a dry weight basis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


