[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Always Available Construction LLC.
Project Name: Yard Soil Cleanup
Project # N/A
Order ID # Q3157
Test Name: TCLP BNA
A. Number of Samples and Date of Receipt:

2 Solid samples were received on 09/19/2025.

1 Water sample was received on 09/19/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

TCLP BNA. This data package contains results for TCLP BNA.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_G using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of TCLP BNA was based on method 8270E and extraction was done based on method 3510 and TCLP extraction method was 1311.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for VNJ-238MS [2,4,6-Tribromophenol - 147%], VNJ-238MSD [2,4,6-Tribromophenol - 151%], 20-DEAD-1 [2,4 and6-Tribromophenol - 156%], PB169754TB [2,4 and6-Tribromophenol - 138%]. AS per method one Acid surrogate is allowed to fail, Therefore no corrective action required.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS {Q3133-02MS} with File ID: BG064423.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[124%], 2,4,6-Trichlorophenol[116%] and Hexachlorobenzene[124%] due to matrix interference.
The MSD {Q3133-02MSD} with File ID: BG064424.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[124%], 2,4,6-Trichlorophenol[124%] and Hexachlorobenzene[124%] due to matrix interference.
The RPD were met for all analysis.
The Blank Spike for {PB169783BS} with File ID: BG064420.D met requirements for all compounds except for Hexachlorobutadiene[103%]. But associaed samples have not positive hit for this compound therefore no corrective action was taken.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID BG064418.D met the requirements except for 1,4-Dichlorobenzene,2,4,5-Trichlorophenol,2,4,6-Trichlorophenol,2,4-Dinitrotoluene, Hexachlorobenzene, Hexachlorobutadiene,Pentachlorophenol,2,4,6-Tribromophenol,2-Fluorophenol,Nitrobenzene-d5 and Terphenyl-d14. But associaed samples have not positive hit for these compounds therefore no corrective action was taken.
The Tuning criteria met requirements.          

E. Additional Comments:
The temperature of the samples at the time of receipt was 26.1°C.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


