[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Always Available Construction LLC.
Project Name: Yard Soil Cleanup
Project # N/A
Order ID # Q3157
Test Name: VOC-TCLVOA-10
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 09/19/2025.

1 Water sample was received on 09/19/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOC-TCLVOA-10. This data package contains results for VOC-TCLVOA-10.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis performed on instrument MSVOA_W were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-TCLVOA-10 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD for {VW0922SBSD01} with File ID: VW032226.D met criteria except for Trichlorofluoromethane[34%]. Due to difference in BS and BSD concentrations. 
The Blank Spike Duplicate for {VN0924WBSD01} with File ID: VN087937.D met requirements for all compounds except for 1,2-Dibromo-3-Chloropropane[116%], 1,3-Dichlorobenzene[110%] and 1,4-Dichlorobenzene[108%]. Failing high  but associated samples have not positive hit for these compounds therefore no corrective action was taken.

The Blank Spike for {VW0922SBS01} with File ID: VW032225.D met requirements for all compounds except for Chloroethane[132%].Failing high  but associated samples have not positive hit for this compound therefore no corrective action was taken.

The Blank Spike Duplicate for {VW0922SBSD01} with File ID: VW032226.D met requirements for all compounds except for 1,1,2-Trichlorotrifluoroethane[136%], 1,1-Dichloroethene[133%], Chloroethane[138%] and Trichlorofluoromethane[138%].Failing high  but associated samples have not positive hit for these compounds therefore no corrective action was taken.
The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Initial Calibration method (82W082625S.M) for Methylene chloride, this compound is passing on Quadratic regression.
The Continuous Calibration File ID VN087933.D met the requirements except for 2-Butanone,2-Hexanone and Acetone. Failing high  but associated samples have not positive hit for these compounds therefore no corrective action was taken.
The Tuning criteria met requirements.          

E. Additional Comments:
Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

The temperature of the samples at the time of receipt was 26.1°C.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


