[image: image1.png]AII|anoe

ECHNICA






284 Sheffield Street, Mountainside, NJ 7092, Phone: 908 789 8900, Fax: 908 789 8922


CASE NARRATIVE

G Environmental
Project Name: Amsterdam
Project # N/A
Order ID # Q3172
Test Name: VOCMS Group1,Mercury,Metals ICP-TAL,Anions Group1
A. Number of Samples and Date of Receipt:

3 Water samples were received on 09/23/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group1,Mercury,Metals ICP-TAL,Anions Group1. This data package contains results for VOCMS Group1(8260-Low),Mercury(7470A),Metals ICP-TAL(6010D),Anions Group1(300.0).

C. Analytical Techniques:

VOCMS Group1 : The analysis performed on instrument MSVOA_V were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UI. The analysis of VOCMS Group1 was based on method 8260D.

Mercury,Metals ICP-TAL : The analysis of Metals ICP-TAL was based on method 6010D, digestion based on method 3010 (waters). The analysis and digestion of Mercury was based on method 7470A.

Wetchem : The analysis of Anions Group1 was based on method 300.0 and extraction was done based on method 8015B.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
Sample FW7D was diluted due to high concentrations for Sulfate.
The Surrogate recoveries were met for all analysis except following 

VOCMS Group1 : FW7D [Toluene-d8 - 83%], FW7DDL [Toluene-d8 - 80%], due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported.

The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS recoveries met the requirements for all compounds except following Mercury,Metals ICP-TAL : The Matrix Spike (TOTES COMP#1MS) analysis met criteria for all compounds except for Sodium and Zinc due to Chemical Interference during Digestion Process.

The MSD recoveries met the requirements for all compounds except following 

Mercury,Metals ICP-TAL : The Matrix Spike Duplicate (TOTES COMP#1MSD) analysis met criteria for all compounds except for Sodium due to Chemical Interference during Digestion Process.

The RPD were met for all analysis except following 

VOCMS Group1 : The RPD for {VV0924WBSD02} with File ID: VV039165.D met criteria except for Acetone[32%], Bromomethane[34%] due to difference in results of BS and BSD.

The Blank Spike met requirements for all compounds except following 

VOCMS Group1 : The Blank Spike for {VV0924WBS02} with File ID: VV039150.D met requirements for all compounds except for 4-Methyl-2-Pentanone[120%] is failing high but no positive hit in associate sample therefore no corrective action taken.

The Blank Spike Duplicate met requirements for all compounds

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements except following 

VOCMS Group1 : The %RSD is greater than 20% in the Initial Calibration method (82V092225W.M) for 2-Hexanone passing on Linear regression and Chloroethane,Acetone ,Methylene Chloride,Styrene,1,2-Dibromo-3-Chloropropane are passing on Quadtratic Regression.

The Continuous Calibration met the requirements except following 

VOCMS Group1 : The Continuous Calibration File ID VV039146.D met the requirements except for Dichlorodifluoromethane is failing marginally low therefore no corrective action taken.
The Continuous Calibration File ID VV039168.D met the requirements except for Bromomethane is failing marginally low therefore no corrective action taken.
The Tuning criteria met requirements.          

The Duplicate analysis met criteria for all samples.
The Serial Dilution met the acceptable requirements.
E. Additional Comments:
VOCMS Group1 : Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


