[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Weston Solutions, Inc.
Project Name: RFP 916
Project # N/A
Order ID # Q3177
Test Name: SPLP VOA
A. Number of Samples and Date of Receipt:

3 Solid samples were received on 09/23/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SPLP VOA. This data package contains results for SPLP VOA.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of SPLP VOA was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS {Q3179-02MS} with File ID: VN087957.D recoveries met the requirements for all compounds except for 2-Hexanone[132%], Methylene Chloride[116%] due to matrix interference.
The MSD {Q3179-03MSD} with File ID: VN087958.D recoveries met the requirements for all compounds except for 2-Chloroethyl vinyl ether[132%], Methyl iodide [123%] due to matrix interference.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID VN087948.D met the requirements except for 2-Chloroethyl Vinyl ether is failing high but no positive hit in associate sample therefore no corrective action taken.
The Tuning criteria met requirements.          

E. Calculation:

Water Calculation in ug/L

(A x )(I s ) (Df)

(Ais) (RRF)(V0)

Where,

Ax = Area for the compound to be measured

Ais = Area for the specific internal standard

Is = Amount of internal standard added in nanograms (ng)

RRF = Relative response factor of the initial calibration curve standard.

Vo = Volume of water purged in milliliters (mL)

Df = Dilution factor.

F. Additional Comments:
Samples for MS/MSD for VOC analysis were not provided with this set of samples therefore lab used from another project.

Trip Blank was not provided with this set of samples.
G. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


