          [image: image2.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
          [image: image1.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3178

	MATRIX: Solid

	METHOD: 8081B/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.


	
	
	(

	2.
	Standard Summary Submitted.


	
	
	(

	3.
	Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD.

The Initial Calibration met the requirements.

The Continuous Calibration File ID PL097449.D met the requirements except for Endrin  is failing in 2nd column but passing in first column therefore no corrective action taken.

 ASK  wdcalcomm " "  \* MERGEFORMAT 

	
	
	(

	 4.
	Blank Contamination - If yes, list compounds and concentrations in each blank:



	
	(
	

	 5.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for EME-TS14-01MS [Decachlorobiphenyl(1)13%, Decachlorobiphenyl(2)15%] but original sample and MSD passing for surrogate therefore no corrective action taken.

	
	(
	

	 6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3178-02MS} with File ID: PL097444.D recoveries met the requirements for all compounds except for [4,4-DDE(1)52% - 4,4-DDE(2)49%], [4,4-DDT(1)44% - 4,4-DDT(2)31%], [Aldrin(1)48% - Aldrin(2)51%], [alpha-BHC(1)55% - alpha-BHC(2)57%], [Endosulfan I(1)54% - Endosulfan I(2)45%], [Endosulfan sulfate(1)45% - Endosulfan sulfate(2)33%], [Endrin aldehyde(1)27% - Endrin aldehyde(2)20%], [Endrin ketone(1)52% - Endrin ketone(2)36%], [Endrin(1)50% - Endrin(2)45%], [gamma-BHC (Lindane)(1)41% - gamma-BHC (Lindane)(2)43%], [Heptachlor(1)39% - Heptachlor(2)40%], [Methoxychlor(1)41% - Methoxychlor(2)30%] due to matrix interference.

The MSD {Q3178-03MSD} with File ID: PL097445.D recoveries met the requirements for all compounds except for [4,4-DDT(1)50% - 4,4-DDT(2)40%], [Endosulfan sulfate(1)54% - Endosulfan sulfate(2)45%], [Endrin aldehyde(1)27% - Endrin aldehyde(2)20%], [gamma-BHC (Lindane)(1)18% - gamma-BHC (Lindane)(2)17%], [Heptachlor(1)28% - Heptachlor(2)27%], [Methoxychlor(1)46% - Methoxychlor(2)41%] and due to matrix interference.

The Blank Spike met requirements for all compounds.

The RPD for {Q3178-03MSD} with File ID: PL097445.D met criteria except for 4,4-DDD[31%], 4,4-DDE[27%], Aldrin[27%], alpha-BHC[24%], alpha-Chlordane[23%], beta-BHC[21%], delta-BHC[29%], Dieldrin[26%], Endosulfan II[21%], Endrin ketone[21%], Endrin[23%], gamma-BHC (Lindane)[78%], Heptachlor epoxide[23%] and Heptachlor[33%] due to difference in results of MS and MSD.

	
	(
	

	7.
	Retention Time Shift Meet Criteria (if applicable)

Comments: 


	
	
	(

	8.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: 
The soil samples  results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date
[image: image1.png][image: image2.png]