ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS VOA CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3201

	MATRIX: Water

	METHOD: 8260-Low

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS  Tuning Specifications

BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY 

ASP CLP, CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.
	
	(
	

	
	The %RSD is greater than 20% in the Initial Calibration method (82V092225W.M) for 2-Hexanone passing on Linear regression and Chloroethane, Acetone ,Methylene Chloride, Styrene, 1,2-Dibromo-3-Chloropropane are passing on Quadtratic Regression.

The Continuous Calibration File ID VV039242.D met the requirements except for Bromoform  is failing high but no positive hit in associate sample while Ethyl Benzene, m/p-Xylenes, o-Xylene, Toluene these compounds failing high associated sample required dilution due to high concentration of compounds Therefore, sample was reanalyzed with dilution and reported and 4-Bromofluorobenzene which is not our target compound therefore no corrective action taken.

The Continuous Calibration File ID VV039265.D met the requirements except for Bromomethane, Carbon Disulfide, Chloromethane, Cyclohexane, Dichlorodifluoromethane is failing low but only dilution sample analyzed under this CCAL and dilution not required for mentioned analyte while 1,2-Dichloroethane-d4 and Toluene-d8 which is not our target compound therefore no corrective action taken.

The Continuous Calibration File ID VX048009.D met the requirements except for Tert butyl alcohol is failing marginally low therefore no corrective action taken.
	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Samples MW1, MW3 was diluted due to bad matrix.

Samples MW6, MW14 and MW1 were diluted due to high concentrations.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.
_______________________________________



__________________________

QA REVIEW







Date

