[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Yannuzzi Group, Inc.
Project Name: Edison Yard
Project # N/A
Order ID # Q3238
Test Name: EPH,PCB,Pesticide-TCL,TPH GC
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 09/29/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for EPH(NJEPH), PCB(8082A), Pesticide-TCL(8081B), TPH GC(8015D).

C. Analytical Techniques:

PCB : The analyses were performed on instrument GCECD_P. The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analysis of PCBs was based on method 8082A and extraction was done based on method 3541.

TPH GC : The analysis were performed on instrument FID_F. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 13302.The analysis of TPH GC was based on method 8015D and extraction was done based on method 3541.

Pesticide-TCL : The analysis was performed on instrument ECD_L.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis of Pesticide-TCLs was based on method 8081B and extraction was done based on method 3541.

EPH : The analyses were performed on instrument FID_D. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis were performed on instrument FID_E. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis of EPHs was based on method NJEPH and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis. 

The Surrogate recoveries were met for all analysis. 
The Retention Times were met for all analysis. 

The MS recoveries met the requirements for all compounds except following 

PCB : The MS recoveries for {Q3240-01MS} with File ID: PP075446.D met requirements for all samples except for [AR1016(1)176% - AR1016(2)150%]these compounds did not meet the NJDKQP criteria and in-house criteria due to matrix interference.

The MSD recoveries met the requirements for all compounds except following 

PCB : The MSD {Q3240-01MSD} with File ID: PP075447.D recoveries met requirements for all samples except for [AR1016(1)165%]this compound did not meet the NJDKQP criteria and in-house criteria while, [AR1016(2)144%] this compound did not meet the NJDKQP criteria but met the in-house criteria due to matrix interference.

The RPD were met for all analysis. 

The Blank Spike met requirements for all compounds.The Blank Spike Duplicate met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination. 

The Initial Calibration met the requirements. 

The Continuous Calibration met the requirements except following 

PCB : The Continuous Calibration File ID PP075434.D met the requirements except for Decachlorobiphenyl is failing in 2nd column however it is passing in 1st column therefore no corrective action taken.

E. Additional Comments:
PCB : The temperature of the samples at the time of receipt was 20.5°C.

The soil samples  results  are based on a dry weight basis.

TPH GC : The temperature of the samples at the time of receipt was 20.5°C.

The soil samples  results  are based on a dry weight basis.

Pesticide-TCL : The soil samples  results  are based on a dry weight basis.

The temperature of the samples at the time of receipt was 20.5°C.

EPH : The temperature of the samples at the time of receipt was 20.5°C.

The soil samples  results  are based on a dry weight basis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


