[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

CDM Smith
Project Name: MWCO CJO WTP Edison 295110
Project # N/A
Order ID # Q3266
Test Name: EPH,Herbicide,PCB,Pesticide-TCL
A. Number of Samples and Date of Receipt:

2 Solid samples were received on 10/01/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for EPH(NJEPH), Herbicide(8151A), PCB(8082A), Pesticide-TCL(8081B).

C. Analytical Techniques:

Pesticide-TCL : The analysis was performed on instrument ECD_L.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis of Pesticide-TCLs was based on method 8081B and extraction was done based on method 3541.

PCB : The analyses were performed on instrument GCECD_P. The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analyses were performed on instrument GCECD_O. The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analysis of PCBs was based on method 8082A and extraction was done based on method 3541.

Herbicide : The analysis was performed on instrument ECD_S. The front column is RTX-CLPesticides which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 11139. The rear column is RTX-CLPesticides2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 11324. The analysis of Herbicides was based on method 8151A and extraction was done based on method 3541.

EPH : The analyses were performed on instrument FID_D. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis were performed on instrument FID_E. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis of EPHs was based on method NJEPH and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis. 

The Surrogate recoveries were met for all analysis. 
The Retention Times were met for all analysis. 

The MS recoveries met the requirements for all compounds except following 

Pesticide-TCL : The MS {Q3266-01MS} with File ID: PL097503.D recoveries met the requirements for all compounds except for [Heptachlor(2)- 60%], [Methoxychlor(2)- 53%] due to matrix interference.

Herbicide : The MS {Q3266-01MS} with File ID: PS031959.D recoveries met the requirements for all compounds except for [Dinoseb(1)0% - Dinoseb(2)0%]  due to matrix interfernce.

The MSD recoveries met the requirements for all compounds except following 

Pesticide-TCL : The MSD {Q3266-01MSD} with File ID: PL097504.D recoveries met the requirements for all compounds except for [Heptachlor(2)- 61%] due to matrix interference.

Herbicide : The MSD {Q3266-01MSD} with File ID: PS031960.D recoveries met the requirements for all compounds except for [Dinoseb(1)0% - Dinoseb(2)0%] due to matrix interfernce.

The RPD were met for all analysis. 

The Blank Spike met requirements for all compounds.The Blank Spike Duplicate met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination. 

The Initial Calibration met the requirements. 

The Continuous Calibration met the requirements except following 

PCB : The Continuous Calibration File ID PO114211.D met the requirements except for Aroclor-1260(Peak-04),Aroclor-1260(Peak-05) is failing in 1st column, however it is passed in 2nd column therefore no corrective action was taken.

E. Additional Comments:
Pesticide-TCL : The soil samples  results  are based on a dry weight basis.

PCB : The soil samples  results  are based on a dry weight basis.

Herbicide : The soil samples  results  are based on a dry weight basis.

EPH : The soil samples  results  are based on a dry weight basis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


