ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3266

	MATRIX: Solid

	METHOD: 8270E/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BF092325.M for 2-Nitrophenol & 2,4-Dinitrotoluene passing on Linear regression,  and 2,4-Dinitrophenol, 4,6-Dinitro-2-methylphenol, they  are passing on Quadratic regression.

The %RSD is greater than 20% in the Initial Calibration (Method 8270-BF100625.M)  for 2,4-Dinitrophenol  this  Compounds is passing on Linear regression.

The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP091225.M)  for 2,4-Dinitrophenol & 4-Nitrophenol.  These Compounds are passing on Linear regression.

The Continuous Calibration File ID BF143853.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2,6-Dinitrotoluene,4-Bromophenyl-phenylether,Hexachlorobenzene,Pentachlorophenol and 2,4,6-Tribromophenol. There is no positive hit in associated samples, therefor no further corrective action was taken.

The Continuous Calibration File ID BP025924.D met the requirements except for Benzaldehyde.  There is no positive hit in associated samples, therefor no further corrective action was taken


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for, 
ENV-3-6FT [Terphenyl-d14 - 116%],As per SOW one base surrogate allowed fail, Therefor no further corrective action was taken. 
And,

ENV-3-6FTMS [Nitrobenzene-d5 - 120%, Phenol-d6 - 115%, Terphenyl-d14 - 139%] and ENV-3-6FTMSD [Nitrobenzene-d5 - 112%, Terphenyl-d14 - 127%]. Failed same surrogate in MS/MSD due to matrix interference, Therefor no further corrective action was taken. 

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3266-01MS} with File ID: BF143864.D recoveries met the requirements for all compounds except for 2,6-Dinitrotoluene[132%], Recovery failed due to matrix interference, Therefor no further corrective action was taken.

The Blank Spike for {PB169989BS} with File ID: BP025926.D met requirements for all compounds except for 1,2,4,5-Tetrachlorobenzene[106%]. there is no positive hit for failed compound in associated samples, Therefor no further corrective action was taken.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
The soil samples results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date

