ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3267

	MATRIX: Water

	METHOD: 200.7

	


	
	
	NA
	NO
	YES

	   1.
	Calibration Summary met criteria.


	
	
	(

	   2.
	ICP Interference Check Sample Results Summary Submitted.


	
	
	(

	   3.
	Serial Dilution Summary (if applicable) Submitted.


	
	
	(

	   4.
	Laboratory Control Sample Summary (if applicable) Submitted.


	
	
	(

	   5.
	Blank Contamination - If yes, list compounds and concentrations in each blank:



	
	(
	

	   6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	   7.
	Sample Duplicate Analysis Met QC Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	   8.
	Digestion Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	   9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: Sample Q3267-01 was analyzed with dilution x50 as straight sample because of very high Sodium concentration in the sample, the previous analytical run which was analyzed on 10/03/25, all QC (CCV,CCB) fails due to suppression effect of this Sodium salt and carryover was extremely high in CCB, so that data was not reported because of QC failure So sample was analyze in this run as straight x50 and x100 dilutions and was reported from x50 dilution.
_______________________________________



__________________________

QA REVIEW







Date


