[image: image1.png]AII|anoe

ECHNICA






284 Sheffield Street, Mountainside, NJ 7092, Phone: 908 789 8900, Fax: 908 789 8922


CASE NARRATIVE

G Environmental
Project Name: DPW
Project # N/A
Order ID # Q3278
Test Name: VOCMS Group2,Metals Group3,Anions Group1
A. Number of Samples and Date of Receipt:

1 Water sample was received on 10/02/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group2,Metals Group3,Anions Group1. This data package contains results for VOCMS Group2(8260-Low),Metals Group3(6010D),Anions Group1(300.0).

C. Analytical Techniques:

VOCMS Group2 : The analysis performed on instrument MSVOA_X were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UI. The analysis of VOCMS Group2 was based on method 8260-Low.

Metals Group3 : The analysis of Metals Group3 was based on method 6010D and digestion based on method 3010 (waters).

Wetchem : The analysis of Anions Group1 was based on method 300.0.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
Metals Group3 : Sample MW10 was diluted due to high concentrations for Manganese.

The Surrogate recoveries were met for all analysis.
The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the acceptable requirements.
The RPD recoveries met criteria.
The Blank Spike met requirements for all compounds.  
The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          
The Duplicate analysis met criteria for all compounds.
The Serial Dilution met the acceptable requirements.
E. Additional Comments:
Metals Group3 : In analytical sequence LB137466, The Result was outside of acceptance limit for Sodium of CCB05 and CCB06 but, no any sample associated under these CCBs.

VOCMS Group2 : Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


