ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3298

	MATRIX: Water

	METHOD: 8270-Modified/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The Initial Calibration met the requirements.

The Continuous Calibration File ID BN038066.D met the requirements except for Nitrobenzene-d5 and Phenol-d6, The failure compounds not associated with the client parameters list, therefore no corrective action was taken. 

The Continuous Calibration File ID BN038083.D met the requirements except for 2,4,6-Tribromophenol and Phenol-d6, The failure compounds not associated with the client parameters list, therefore no corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for BP-VPB-184-EB-20251003 [2-Fluorobiphenyl - 110%, Terphenyl-d14 - 144%], BP-VPB-184-GW-740-742MS [2-Fluorobiphenyl - 183%, Terphenyl-d14 - 269%] and BP-VPB-184-GW-740-742MSD [2-Fluorobiphenyl - 218%, Terphenyl-d14 - 553%], The failure surrogates  not associated with the client parameters list, therefore no corrective action was taken, and VPB184-HYD-20251003 [ 2-Fluorobiphenyl - 117%, 2-Methylnaphthalene-d10 - 0%, Fluoranthene-d10 - 14%,  Terphenyl-d14 - 147%], due to matrix interference. 

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3298-06MS} with File ID: BN038086.D recoveries met the requirements for all compounds except for 1,4-Dioxane[168%] due to matrix interference.

The MSD {Q3298-07MSD} with File ID: BN038087.D recoveries met the requirements for all compounds except for 1,4-Dioxane[173%] due to matrix interference.

The Blank Spike met requirements for all compounds.


	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: The Internal Standards Areas were met for all analysis except for VPB184-HYD-20251003, BP-VPB-184-GW-740-742MS and BP-VPB-184-GW-740-742MSD, The failure Internal Standard not associated with the client parameters list, therefore no corrective action was taken.

	
	(
	

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS:
The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).” 
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.
_______________________________________



__________________________

QA REVIEW







Date

