[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Ronald Adams
Project Name: 19 Ashlynn Court
Project # N/A
Order ID # Q3301
Test Name: EPH,Gasoline Range Organics
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 10/06/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for EPH(NJEPH), Gasoline Range Organics(8015D).

C. Analytical Techniques:

EPH : The analysis were performed on instrument FID_C. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analyses were performed on instrument FID_D. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis of EPHs was based on method NJEPH and extraction was done based on method 3541.

Gasoline Range Organics : The analysis performed on instrument FID_B were done using GC column RTX502.2 which is 60 meters, 0.53mm ID, 3.0 um df, cat#10909.The analysis of Gasoline Range Organics was based on method 8015D.
D. QA/ QC Samples:

The Holding Times were met for all analysis except following 
EPH : The Holding Times were met for all analysis except for SOIL. As sample received late in Lab.

Gasoline Range Organics : The Holding Times were met for all analysis except for SOIL. As sample received late in Lab
The Surrogate recoveries were met for all analysis except following 

EPH : 

SOIL Aliphatic [1-chlorooctadecane (SURR) - 191%],

SOIL Aromatic [2-Bromonaphthalene (SURR) - 190%,  2-Flurobiphenyl (SURR) - 144%],

SOILDL Aliphatic [1-chlorooctadecane (SURR) - 164%], 

SOILDL Aromatic [2-Bromonaphthalene (SURR) - 181% , 2-Flurobiphenyl (SURR) - 150%], 

SOILDL2 Aliphatic [1-chlorooctadecane (SURR) - 164%], 

SOILDL2 Aromatic [2-Bromonaphthalene (SURR) - 360% , 2-Flurobiphenyl (SURR) - 478%, ortho-Terphenyl (SURR) - 258%], 

but this sample was required further dilution as well due to high concentration, 

therefore original and Dilution analysis were reported and no further corrective action taken.

SOILMS Aromatic [2-Bromonaphthalene (SURR) - 184%],

SOILMSD Aromatic [2-Bromonaphthalene (SURR) - 180% , 2-Flurobiphenyl (SURR) - 144%], MSMSD confirm with its original sample.
The Retention Times were met for all analysis. 

The MS recoveries met the requirements for all compounds except following 

EPH : The MS {Q3301-01MS} with File ID: FC069965.D recoveries met the requirements for all compounds except for Aliphatic [ n-Hexadecane (C16)- 19%], [n-Octadecane (C18) - 13%], [n-Hexatriacontane (C36) - 148%], [n-Octatriacontane (C38) -149%] due to matrix interference. & for [Naphthalene (C11.7)- 24%, 2-methylnaphthalene (C12.89)- 250%], these analytes are only being monitoring in aliphatic fraction.

The MS {Q3301-01MS} with File ID: FD049834.D recoveries met the requirements for all compounds except for Aromatic [Acenaphthylene (C15.06) - 143%], [Acenaphthene (C15.5) - 285%], Flouorene (C16.55) - 148%], [Anthracene (C19.43) - 497%], [Pyrene (C20.8) - 178%], [Fluoranthene (C21.85) - 186%], [Benzo[a]anthracene (C26.37) - 205%], [Chrysene (C27.41) - 414%], [Bnezo[k]fluoranthene (C30.14) - 410%],[Dibenz[a,h]anthracene (C30.36) - 381%], [benzo[b]fluoranthene (C30.41) - 197%], [Benzo[a]pyrene (C31.34) - 191%], [Benzo[g,h,i]perylene (C34.01) - 190%], [Indeno[1,2,3-cd]pyrene (C35.01) - 216%] due to matrix interference.

The MSD recoveries met the requirements for all compounds except following 

EPH : The MSD {Q3301-01MSD} with File ID: FC069966.D recoveries met the requirements for all compounds except for Aliphatic [ n-Decane (C10)- 146%], [ n-Tetradecane (C14)- 155%], [n-Octadecane (C18) - 479%], [n-Octatriacontane (C38) -150%], [n-Tetracontane (C40) - 142%] due to matrix interference. & for [Naphthalene (C11.7)- 39%, 2-methylnaphthalene (C12.89)- 519%], these analytes are only being monitoring in aliphatic fraction.

The MSD {Q3301-01MSD} with File ID: FD049835.D recoveries met the requirements for all compounds except for Aromatic [Acenaphthene (C15.5) - 282%], Flouorene (C16.55) - 142%], [Anthracene (C19.43) - 494%], [Pyrene (C20.8) - 177%], [Fluoranthene (C21.85) - 181%], [Benzo[a]anthracene (C26.37) - 202%], [Chrysene (C27.41) - 409%], [Bnezo[k]fluoranthene (C30.14) - 408%],[Dibenz[a,h]anthracene (C30.36) - 380%], [benzo[b]fluoranthene (C30.41) - 195%], [Benzo[a]pyrene (C31.34) - 190%], [Benzo[g,h,i]perylene (C34.01) - 189%], [Indeno[1,2,3-cd]pyrene (C35.01) - 213%] due to matrix interference.

The RPD were met for all analysis except following 

EPH : The RPD for {Q3301-01MSD} with File ID: FC069966.D met criteria except for Aliphatic  [2-methylnaphthalene (C12.89) - 69.96%], [n-Hexadecane (C16) - 142.42%], [n-Octadecane (C18) - 189.43%], [n-Eicosane (C20) - 71.35%] due to difference in MSMSD concentrations.

The Blank Spike met requirements for all compounds.
The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination. 
The Initial Calibration met the requirements. 
The Continuous Calibration met the requirements. 

Gasoline Range Organics : Samples SOIL was diluted due to bad matrix.

EPH : Sample SOIL,SOILDL were diluted due to high concentration for Aliphatic and Aromatic compounds. 
E. Additional Comments:
The soil samples  results  are based on a dry weight basis.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


