[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Weston Solutions, Inc.
Project Name: RFP 918
Project # N/A
Order ID # Q3321
Test Name: TCLP BNA
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 10/09/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

TCLP BNA. This data package contains results for TCLP BNA.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_G using GC Column ZB-Semi Volatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of TCLP BNA was based on method 8270E and extraction was done based on method 3510 and TCLP extraction method was 1311.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis. 
The MS {Q3321-02MS} with File ID: BG064503.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[120%], 2,4,6-Trichlorophenol[118%] and Hexachlorobenzene[116%], due to matrix interference due to matrix interference.
The MSD {Q3321-02MSD} with File ID: BG064504.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[116%], due to matrix interference due to matrix interference.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
F. Manual Integration Comments:
Concentration of Water Sample:

Concentration ug/L = (Ax) (Is) (Vt) (DF) (GPC)            





__                                      

                                   (Ais) (RRF) (Vo) (Vi)

Where,

Ax = Area of the characteristic ion for the compound to be measured.

Ais = Area of the characteristic ion for the internal standard.

Is = Amount of internal standard injected in ng.

Vo = Volume of water extracted in mL.

Vi = Volume of extract injected in uL.

Vt = Volume of the concentrated extract in uL 

____

RRF = Mean Relative Response Factor determined from the initial calibration standard.

GPC = Vin  = GPC factor (If no GPC is performed, GPC=1)

            Vout

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


