[image: image1.png]AII|anoe

ECHNICA






284 Sheffield Street, Mountainside, NJ 7092, Phone: 908 789 8900, Fax: 908 789 8922


CASE NARRATIVE

Aramark Uniforms
Project Name: Monthly 2025
Project # N/A
Order ID # Q3368
Test Name: TCLP VOA,TCLP BNA Group1,PCB,TCLP Herbicide,TCLP Pesticide,TCLP Mercury,TCLPMetals Group2, ASTM Ammonia,ASTM COD,ASTM Oil and Grease,ASTM TS,Corrosivity,Ignitability,Oil and Grease,Paint Filter,pH,Reactive Cyanide,Reactive Sulfide,TS,TVS
A. Number of Samples and Date of Receipt:

4 Solid samples were received on 10/16/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

TCLP VOA,TCLP BNA Group1,PCB,TCLP Herbicide,TCLP Pesticide,TCLP Mercury,TCLPMetals Group2, ASTM Ammonia,ASTM COD,ASTM Oil and Grease,ASTM TS, Corrosivity, Ignitability,Oil and Grease,Paint Filter,pH,Reactive Cyanide,Reactive Sulfide,TS,TVS This data package contains results for TCLP VOA(8260D),TCLP BNA Group1(8270E),PCB(8082A),TCLP Herbicide(8151A),TCLP Pesticide(8081B),TCLP Mercury(7470A),TCLPMetals Group2(6010D), ASTM Ammonia(SM4500-NH3),ASTM COD(SM5220 D),ASTM Oil and Grease(1664A),ASTM TS(SM2540 B),Corrosivity(9045D),Ignitability(1030),Oil and Grease(9071B),Paint Filter(9095B),pH(9045D),Reactive Cyanide(9012B),Reactive Sulfide(9034),TS(SM2540 B),TVS(160.4).

C. Analytical Techniques:

TCLP VOA : The analysis performed on instrument MSVOA_X were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UI. The analysis of TCLP VOA was based on method 8260D and TCLP extraction method was 1311.

TCLP BNA Group1 : The samples were analyzed on instrument BNA_P using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of TCLP BNA Group1 was based on method 8270E and extraction was done based on method 3510 and TCLP extraction method was 1311.

TCLP Pesticide : The analysis was performed on instrument ECD_D.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis of TCLP Pesticides was based on method 8081B and extraction was done based on method 3510 and TCLP extraction method was 1311.

TCLP Herbicide : The analysis was performed on instrument ECD_S. The front column is RTX-CLPesticides which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 11139. The rear column is RTX-CLPesticides2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 11324. The analysis of TCLP Herbicides was based on method 8151A and extraction was done based on method 3510 and TCLP extraction method was 1311.

PCB : The analyses were performed on instrument GCECD_Q.The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analyses were performed on instrument GCECD_O. The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analysis of PCBs was based on method 8082A and extraction was done based on method 3541.

TCLP Mercury,TCLPMetals Group2 : The analysis of TCLPMetals Group2 was based on method 6010D, digestion based on method 3010 (waters). The analysis and digestion of TCLP Mercury was based on method 7470A and TCLP extraction method was 1311.

Wetchem : The analysis of Ignitability was based on method 1030, The analysis of TVS was based on method 160.4, The analysis of ASTM Oil and Grease was based on method 1664A, The analysis of Reactive Cyanide was based on method 9012B, The analysis of Reactive Sulfide was based on method 9034, The analysis of pH was based on method 9040C, The analysis of Corrosivity,pH was based on method 9045D, The analysis of Oil and Grease was based on method 9071B, The analysis of Paint Filter was based on method 9095B, The analysis of ASTM TS,TS was based on method SM2540 B, The analysis of ASTM Ammonia was based on method SM4500-NH3 and The analysis of ASTM COD was based on method SM5220 D
D. QA/ QC Samples:

The Holding Times were met for all analysis except following Wetchem : OUTSIDE-DUMPSTER of Corrosivity, pH as sample was receive out of holding time. for PRESS-SLUDGE of Corrosivity, pH as sample was receive out of holding time. OUTSIDE-DUMPSTER of TS, TVS as sample was analyzed out of holding time due to activated sample. for PRESS-SLUDGE of TS, TVS as sample was analyzed out of holding time for activated sample due to samples activate later. TCLP Herbicide: TCLP extraction to normal extraction out of hold because sample's activate later.TCLP Pesticide: TCLP extraction to normal extraction out of hold because sample's activate later.

The Surrogate recoveries were met for all analysis except following 

TCLP Pesticide :  OUTSIDE-DUMPSTER [Decachlorobiphenyl(1)38%, Decachlorobiphenyl(2)20%, Tetrachloro-m-xylene(1)46%, Tetrachloro-m-xylene(2)55%], OUTSIDE-DUMPSTERRE [Decachlorobiphenyl(1)44%, Decachlorobiphenyl(2)41%, Tetrachloro-m-xylene(1)48%, Tetrachloro-m-xylene(2)59%], PRESS-SLUDGE [Decachlorobiphenyl(1)27%, Decachlorobiphenyl(2)21%, Tetrachloro-m-xylene(1)15% and Tetrachloro-m-xylene(2)37%], PRESS-SLUDGERE [Decachlorobiphenyl(1)32%, Decachlorobiphenyl(2)28%, Tetrachloro-m-xylene(1)15% and Tetrachloro-m-xylene(2)39%], OUTSIDE-DUMPSTERMS [Decachlorobiphenyl(1)47%, Decachlorobiphenyl(2)31%],  OUTSIDE-DUMPSTERMSD [Decachlorobiphenyl(1)48%, Decachlorobiphenyl(2)35%], due to matrix interference. These samples did not re-extracted as there was no extra volume. Hence samples re-analyzed and original & Reanalysis both are reported.

PCB : PRESS-SLUDGE [Decachlorobiphenyl(1)182%]AS per method one surrogate allowed to fail to meet the criteria per column, No further corrective action was taken.

The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS recoveries met the requirements for all compounds except following 

TCLP Pesticide : The MS {Q3368-06MS} with File ID: PD090945.D recoveries met the requirements for all compounds except for [Endrin(1)68% - Endrin(2)58%], [Heptachlor epoxide(1)72% - Heptachlor epoxide(2)66%] and [Methoxychlor(2)52%] due matrix interference.

TCLP Herbicide : The MS {Q3368-06MS} with File ID: PS032272.D recoveries met the requirements for all compounds except for [2,4,5-TP(Silvex)(1)51%] and [2,4-D(1)48% - 2,4-D(2)58%] due to matrix interference.

TCLP Mercury,TCLPMetals Group2 : The Matrix Spike (TP-6MS) analysis met criteria for all compounds except for Barium due to Chemical Interference during Digestion Process.

 Wetchem : The Matrix Spike (OUTSIDE-DUMPSTERMS) analysis met criteria for all compounds except for Oil and Grease due to sample matrix interference.

The MSD recoveries met the requirements for all compounds except following 

TCLP Pesticide : The MSD {Q3368-06MSD} with File ID: PD090946.D recoveries met the requirements for all compounds except for Endrin(1)72% - Endrin(2)62%], [Heptachlor epoxide(1)76% - Heptachlor epoxide(2)70%] and [Methoxychlor(2)58%] due to matrix interference.

TCLP Herbicide : The MSD {Q3368-06MSD} with File ID: PS032273.D recoveries met the requirements for all compounds except for [2,4,5-TP(Silvex)(1)51%] and [2,4-D(1)48% - 2,4-D(2)58%]due to matrix interference.

TCLP Mercury,TCLPMetals Group2 : The Matrix Spike Duplicate(TP-6MSD) analysis met criteria for all compounds except for Barium due to Chemical Interference during Digestion Process.

 Wetchem : The Matrix Spike Duplicate (OUTSIDE-DUMPSTERMSD) analysis met criteria for all compounds except for Oil and Grease due to sample matrix interference.

The RPD recoveries met criteria.
The Blank Spike met requirements for all compounds.  
The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements except following 

TCLP BNA Group1 : The Continuous Calibration File ID BP026001.D met the requirements except for Terphenyl-d14. Which is not our target compound, therefore no corrective action taken.

TCLP Pesticide : The Continuous Calibration File ID PD090949.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action eas taken. AND Decachlorobiphenyl is failing in 2nd column. failing marginally therefore no corrective action was taken. 

The Continuous Calibration File ID PD090957.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.

The Continuous Calibration File ID PD090961.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.

The Continuous Calibration File ID PD090969.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.

The Tuning criteria met requirements.          

The Duplicate analysis met criteria for all samples.
The Serial Dilution met the acceptable requirements.
TCLP VOA: Samples PRESS-SLUDGE was diluted due to sample has difficult matrix and did not purged well, lab has analyzed sample at 50x straight dilution.

E. Additional Comments:
The soil samples  results  are based on a dry weight basis.

ASTM Ammonia,ASTM COD,ASTM Oil and Grease,ASTM TS,Corrosivity,Ignitability,Oil and Grease,Paint Filter,pH,Reactive Cyanide,Reactive Sulfide,TS,TVS : Sample Q3368-13 and Q3368-14 were analyzed straight dilution for ASTM COD because the original samples were reading over range, only 5X has been reported.

PCB: As per special requirement for this project form-1 and Hit Summary are reported in mg/kg.

TCLP VOA : Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


