[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Aramark Uniforms
Project Name: Monthly 2025
Project # N/A
Order ID # Q3368
Test Name: TCLP Pesticide
A. Number of Samples and Date of Receipt:

4 Solid samples were received on 10/16/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

TCLP Pesticide. This data package contains results for TCLP Pesticide.

C. Analytical Techniques:

The analysis was performed on instrument ECD_D.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis of TCLP Pesticides was based on method 8081B and extraction was done based on method 3510 and TCLP extraction method was 1311.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for  OUTSIDE-DUMPSTER [Decachlorobiphenyl(1)38%, Decachlorobiphenyl(2)20%, Tetrachloro-m-xylene(1)46%, Tetrachloro-m-xylene(2)55%], OUTSIDE-DUMPSTERRE [Decachlorobiphenyl(1)44%, Decachlorobiphenyl(2)41%, Tetrachloro-m-xylene(1)48%, Tetrachloro-m-xylene(2)59%], PRESS-SLUDGE [Decachlorobiphenyl(1)27%, Decachlorobiphenyl(2)21%, Tetrachloro-m-xylene(1)15% and Tetrachloro-m-xylene(2)37%], PRESS-SLUDGERE [Decachlorobiphenyl(1)32%, Decachlorobiphenyl(2)28%, Tetrachloro-m-xylene(1)15% and Tetrachloro-m-xylene(2)39%], OUTSIDE-DUMPSTERMS [Decachlorobiphenyl(1)47%, Decachlorobiphenyl(2)31%],  OUTSIDE-DUMPSTERMSD [Decachlorobiphenyl(1)48%, Decachlorobiphenyl(2)35%], due to matrix interference. These samples did not re-extracted as there was no extra volume. Hence samples re-analyzed and original & Reanalysis both are repotred.
The Retention Times were met for all analysis.
The MS {Q3368-06MS} with File ID: PD090945.D recoveries met the requirements for all compounds except for [Endrin(1)68% - Endrin(2)58%], [Heptachlor epoxide(1)72% - Heptachlor epoxide(2)66%] and [Methoxychlor(2)52%] due matrix interference due to matrix interference.
The MSD {Q3368-06MSD} with File ID: PD090946.D recoveries met the requirements for all compounds except for Endrin(1)72% - Endrin(2)62%], [Heptachlor epoxide(1)76% - Heptachlor epoxide(2)70%] and [Methoxychlor(2)58%] due to matrix interference.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID PD090949.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action eas taken. AND Decachlorobiphenyl is failing in 2nd column. failing marginally therefore no corrective action was taken. 
The Continuous Calibration File ID PD090957.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.

The Continuous Calibration File ID PD090961.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.

The Continuous Calibration File ID PD090969.D met the requirements except for Heptachlor is failing in 1st column, failing high but associated samples have no positive hit for this compund therefore no corrective action was taken.
E. Additional Comments:
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


