ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3399

	MATRIX: Solid

	METHOD: 8270E/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BG102421.M for 2-Nitrophenol,  2-Nitroaniline, 2,6-Dinitrotoluene,   2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenolthese  Compounds are passing on Linear regression,

The Continuous Calibration File ID BG064672.D met the requirements except for 2,4-Dinitrophenol,2,4-Dinitrotoluene,2,6-Dinitrotoluene,2-Nitroaniline,2-Nitrophenol,3-Nitroaniline,4,6-Dinitro-2-methylphenol,Hexachlorocyclopentadiene and Nitrobenzene-d5, The associate samples have no positive hit for these compounds therefore no corrective action was taken.

The Continuous Calibration File ID BG064685.D met the requirements except for 2-Nitrophenol, faling high but The associate samples have no positive hit for these compounds therefore no corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:



	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3399-01MS} with File ID: BG064683.D recoveries met the requirements for all compounds except for 1,4-Dioxane[67%] and Hexachlorocyclopentadiene[149%], due to matrix interference therefore no corrective action was taken.

The MSD {Q3399-01MSD} with File ID: BG064684.D recoveries met the requirements for all compounds except for 1,4-Dioxane[63%], 2-Nitrophenol[126%] and Hexachlorocyclopentadiene[151%], due to matrix interference therefore no corrective action was taken.

The Blank Spike Duplicate met requirements for all compounds.The Blank Spike for {PB170206BS} with File ID: BG064674.D met requirements for all compounds except for 1,4-Dioxane[59%], marginally low, therefore no corrective action was taken, Hexachlorocyclopentadiene[127%] failing high but The associate samples have no positive hit for these compounds therefore no corrective action was taken.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: The Internal Standards Areas were met for all analysis.

	
	(
	

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date

