ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

METALS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3447

	MATRIX: Solid

	METHOD: 6010D,7471B

	


	
	
	NA
	NO
	YES

	   1.
	Calibration Summary met criteria.


	
	
	(

	   2.
	ICP Interference Check Sample Results Summary Submitted.


	
	
	(

	   3.
	Serial Dilution Summary (if applicable) Submitted.

The Serial Dilution (EO-01-102325L) met criteria for all compounds except for Calcium, Chromium, Magnesium, and Manganese due to unknow sample matrix interference.

	
	(
	

	   4.
	Laboratory Control Sample Summary (if applicable) Submitted.


	
	
	(

	   5.
	Blank Contamination - If yes, list compounds and concentrations in each blank:



	
	(
	

	   6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Matrix Spike (EO-01-102325MS) analysis met criteria for all compounds except for Antimony, Beryllium, Selenium, Silver and Vanadium due to Chemical Interference during Digestion process.  The Matrix Spike Duplicate (EO-01-102325MSD) analysis met criteria for all compounds except for Antimony, Barium, Beryllium, Mercury, Selenium and Silver due to Chemical Interference during Digestion process.

	
	(
	

	   7.
	Sample Duplicate Analysis Met QC Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(

	   8.
	Digestion Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	   9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: 
Sample Q3447-01 and Q3447-03 , Silver parameter is Oversaturated so , its Reported from its 5X Dilutions.

The Post Digest Spike (EO-01-102325A) analysis met criteria for all compounds except for Barium, Beryllium, Selenium, Silver and Vanadium due to unknown chemical interference of matrix with the addition of spike amount after digestion and before analysis; matrix has suppression effect during addition of spike.
            In analytical Sequence LB137662, The Result was outside acceptance Limit for Sodium of CCB05 but, no any samples  parameter associated under this CCB.
           In analytical Sequence LB137677, The Result was outside acceptance Limit for Iron of CCB10 but, no any samples parameter associated under this CCB.
           In analytical Sequence LB137662, The Recovery was outside acceptance Limit for Sodium of CCV09 but, no any samples parameter associated under this CCV.

_______________________________________



__________________________

QA REVIEW







Date


