[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

LiRo Engineers, Inc.
Project Name: RFK Bridge RMB-Randall Island
Project # N/A
Order ID # Q3468
Test Name: SVOCMS Group1
A. Number of Samples and Date of Receipt:

6 Water samples were received on 10/24/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

. This data package contains results for SVOCMS Group1(8270E).

C. Analytical Techniques:

SVOCMS Group1 : The samples were analyzed on instrument BNA_G using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOCMS Group1 was based on method 8270E and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis. 

The Surrogate recoveries were met for all analysis. 

The Internal Standards Areas were met for all analysis. 

The Retention Times were met for all analysis. 

The RPD were met for all analysis. 

The Blank Spike met requirements for all compounds.
The Blank Spike Duplicate met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination. 

The Initial Calibration met the requirements. 

The Continuous Calibration met the requirements except following 

SVOCMS Group1 : The Continuous Calibration File ID BG064618.D met the requirements except for Nitrobenzene-d5. Which is not our target compound,  therefore no corrective action taken.

The Continuous Calibration File ID BG064636.D met the requirements except for Nitrobenzene-d5. Which is not our target compound,  therefore no corrective action taken.

The Tuning criteria met requirements. 

E. Additional Comments:
SVOCMS Group1 : The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


