[image: image1.png]AII|anoe

ECHNICA






284 Sheffield Street, Mountainside, NJ 7092, Phone: 908 789 8900, Fax: 908 789 8922


CASE NARRATIVE

Sciacca General Contractors, LLC
Project Name: 7 Rynda Road, Maplewood
Project # N/A
Order ID # Q3474
Test Name: VOC-TCLVOA-10,EPH_F2,TPH GC
A. Number of Samples and Date of Receipt:

7 Solid samples were received on 10/27/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOC-TCLVOA-10,EPH_F2,TPH GC. This data package contains results for VOC-TCLVOA-10(8260D),EPH_F2(NJEPH),TPH GC(8015D).

C. Analytical Techniques:

VOC-TCLVOA-10 : The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-TCLVOA-10 was based on method 8260D.

TPH GC : The analysis were performed on instrument FID_F. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 13302.The analysis of TPH GC was based on method 8015D and extraction was done based on method 3541.

EPH_F2 : The analysis were performed on instrument FID_C. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis were performed on instrument FID_E. The column is RXI-1MS which is 20 meters, 0.18mm ID, 0.18 um df, catalog 10224.The analysis of EPH_F2s was based on method NJEPH and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except following 

EPH_F2 : 5MS [ortho-Terphenyl (SURR) - 38%], 5MSD [ortho-Terphenyl (SURR) - 37%]. Due to sample matrix interference. 

The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the acceptable requirements.
The RPD recoveries met criteria.
The Blank Spike met requirements for all compounds.  
The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements except following 

VOC-TCLVOA-10 : The Continuous Calibration File ID VY023598.D met the requirements except for Ethyl Benzene, Isopropylbenzene and Styrene are failing high but no positive hit in associate sample therefore no corrective action taken.
The Tuning criteria met requirements.       
TPH GC : Samples WASTE was diluted due to bad matrix
E. Additional Comments:
The soil samples  results  are based on a dry weight basis.

VOC-TCLVOA-10 : Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.
Trip Blank was not provided with this set of samples.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


