[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Alliance Technical Group, LLC - Newark
Project Name: NJ Soil PT
Project # N/A
Order ID # Q3483
Test Name: VOCMS Group1
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 10/29/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group1. This data package contains results for VOCMS Group1.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group1 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike for {VY1106SBS01} with File ID: VY023694.D met requirements for all compounds except for Tetrahydrofuran[110%] is failing high but no positive hit in associate sample therefore no corrective action taken.

The Blank Spike Duplicate for {VY1106SBSD01} with File ID: VY023695.D met requirements for all compounds except for Tetrahydrofuran[110%] is failing high but no positive hit in associate sample therefore no corrective action taken.
The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Initial Calibration method (82Y110425S.M) for Naphthalene passing on Linear regression.

The Continuous Calibration File ID VY023692.D met the requirements except for Acrolein.
The Tuning criteria met requirements.          

Sample HW1025-PT-VOA-SOIL was diluted due to high concentration.
E. Additional Comments:
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


